## InstructionsKey
Color Coding
Headings
General Information or Summary Statements that are addressed as specific items
No response needed
Explanatory Statement or notes (not assessable)
Assess as a part of the method
Duplicate Requirement (duplicate citations bold and italicized)
Other citations on same subject matter notes to auditor
Check if any item listed after this statement is a "n"
Remember: | Management system = Quality System

## Vol1Module 2
ID # | Major Topic | Topic | Subject | Category | Citation | Does the laboratory comply with this section? | Yes | No | N/A
1 | Quality Systems Summary |  |  |  | V1M2 1 | Summary of Quality System Requirements - Checklist responses will determine whether or not these have been met. You may use any of these items for assessment if you want.
2 | Quality Systems Summary | Introduction |  |  | V1M2 1 | INTRODUCTION, SCOPE AND APPLICABILITY
3 | Quality Systems Summary | Introduction |  |  | V1M2 1.1 | Introduction
4 | Quality Systems Summary | Introduction | General |  | V1M2 1.1 | Each laboratory shall have a quality system. The laboratory’s quality system is the means by which an organization ensures the quality of the products or services it provides and includes a variety of management, technical, and administrative elements such as:
5 | Quality Systems Summary | Introduction | General | Quality System Components | V1M2 1.1 a) | policies and objectives,
6 | Quality Systems Summary | Introduction | General | Quality System Components | V1M2 1.1 b) | procedures and practices,
7 | Quality Systems Summary | Introduction | General | Quality System Components | V1M2 1.1 c) | organizational authority,
8 | Quality Systems Summary | Introduction | General | Quality System Components | V1M2 1.1 d) | responsibilities, and
9 | Quality Systems Summary | Introduction | General | Quality System Components | V1M2 1.1 e) | performance measures.
10 | Quality Systems Summary | Introduction | General | Use and Documentation | V1M2 1.1 | The quality system provides the framework for planning, implementing, assessing, and improving work performed by an organization so as to provide the client with data of known and documented quality, sufficient to evaluate the usability of the data to the clients needs. The quality system shall be documented in the laboratory’s quality manual and related quality documentation, and shall be referenced in the quality manual.
11 | Quality Systems Summary |  | General | Requirement and Purpose | V1M2 1.1 | This Standard contains detailed quality system requirements for consistent and uniform implementation by the laboratories conducting testing and the consistent and uniform evaluation of those laboratories by accreditation bodies. Each laboratory seeking accreditation under this Standard shall ensure that they are implementing their quality system and that all Quality Control procedures specified in this module are being followed. The Quality Assurance policies, which establish quality control procedures, are applicable to environmental laboratories regardless of size and complexity.
12 | Quality Systems Summary | Introduction | General |  | V1M2 1.1 | This Standard is consistent with ISO/IEC 17025:2005 requirements that are relevant to the scope of environmental testing services.
13 | Quality Systems Summary | Introduction | General | Requirement and Purpose | V1M2 1.1 | All items identified in this document shall be available for an on-site assessment.
14 | Quality Systems Summary | Scope |  |  | V1M2 1.2 | Scope
15 | Quality Systems Summary | Scope |  |  | V1M2 1.2 | The requirements in this document give the basis for a laboratory’s quality system in order to carry out environmental tests. It covers testing performed using reference methods, non-reference methods, and laboratory-developed methods. This document contains the essential elements required to establish a quality system that produces data of known and documented quality, and demonstrates proficiency through the use of proficiency testing and employee training.
16 | Quality Systems Summary | Scope |  |  | V1M2 1.2 | The general requirements of this document apply to all organizations performing environmental tests, regardless of the number of personnel or the degree of environmental testing activities. When the use of the data requires compliance with the Standards, these Standards shall be followed.
17 | Quality Systems Summary | Scope |  |  | V1M2 1.2 | This document is for use by laboratories, clients, regulatory authorities, and accreditation bodies to ensure the laboratory has appropriate management and technical quality systems to perform environmental testing. This document specifies technical, managerial, and documentation requirements needed for assessment by organizations or accreditation bodies to grant approval.
18 | Quality Systems Summary | Scope |  |  | V1M2 1.2 | This document provides the requirements needed for laboratory accreditation. If the requirements of this document are met, the laboratory operates a quality system in conformance with the applicable clauses of ISO/IEC 17025:2005(E).
19 | Definitions | Definitions |  |  | V1M2 3 | Make sure you understand all these terms as they relate to the assessment.
20 | Definitions | Definitions |  |  | V1M2 3 | TERMS AND DEFINITIONS

## V1M2 Management
ID | Major Topic | Topic | Subject | Category | Citation | Does the laboratory comply with this section? | Yes | No | N/A
20 |  |  |  |  | V1M2 4 | MANAGEMENT REQUIREMENTS
21 | 4.1 Organization | 4.1 Organization |  |  | V1M2 4.1 | Organization (ISO/IEC 17025:2005(E), Clause 4.1)
22 | 4.1 Organization | 4.1.1 Legal Status |  | Requirement | V1M2 4.1.1 | The laboratory or the organization of which it is part shall be an entity that can be held legally responsible.
23 | 4.1 Organization | 4.1.2 Laboratory Goals | Responsibility | Requirement | V1M2 4.1.2 | It is the responsibility of the laboratory to carry out its testing and calibration activities in such a way as
24 | 4.1 Organization | 4.1.2 Laboratory Goals | Standard | Requirement | V1M2 4.1.2 | to meet the requirements of this International Standard
25 | 4.1 Organization | 4.1.2 Laboratory Goals | Customers | Requirement | V1M2 4.1.2 | and to satisfy the needs of the customer, the regulatory authorities or organizations providing recognition.
26 | 4.1 Organization | 4.1.3 Management System | Scope | Requirement | V1M2 4.1.3 | The management system shall cover work carried out in the laboratory's permanent facilities, at sites away from its permanent facilities, or in associated temporary or mobile facilities.
27 | 4.1 Organization | 4.1.4 Conflict of Interest | Activities | Requirement | V1M2 4.1.4 | If the laboratory is part of an organization performing activities other than testing and/or calibration, the responsibilities of key personnel in the organization that have an involvement or influence on the testing and/or calibration activities of the laboratory shall be defined in order to identify potential conflicts of interest.
28 | 4.1 Organization | 4.1.4 Conflict of Interest | Activities | Note | V1M2 4.1.4 N1 | NOTE 1: Where a laboratory is part of a larger organization, the organizational arrangements should be such that departments having conflicting interests, such as production, commercial marketing or financing do not adversely influence the laboratory's compliance with the requirements of this International Standard.
29 | 4.1 Organization | 4.1.4 Conflict of Interest | Activities | Note | V1M2 4.1.4 N2 | NOTE 2: If the laboratory wishes to be recognized as a third-party laboratory, it should be able to demonstrate that it is impartial and that it and its personnel are free from any undue commercial, financial and other pressures which might influence their technical judgment. The third-party testing or calibration laboratory should not engage in any activities that may endanger the trust in its independence of judgment and integrity in relation to its testing or calibration activities.
30 | 4.1 Organization | 4.1.5 General Requirements |  |  | V1M2 4.1.5 | The laboratory shall:
31 |  | 4.1.5 General Requirements | Personnel | Requirement | 4.1.5 a) | have managerial and technical personnel who, irrespective of other responsibilities, have the authority and resources needed to carry out their duties, including
32 | 4.1 Organization | 4.1.5 General Requirements | Duties - Management System | Requirement | V1M2 4.1.5 a) | have managerial and technical personnel who have the authority and resources needed to carry out the implementation, maintenance and improvement of the management system,
33 | 4.1 Organization | 4.1.5 General Requirements | Duties- Corrective Actions | Requirement | V1M2 4.1.5 a) | have managerial and technical personnel who have the authority and resources needed to identify the occurrence of departures from the management system or from the procedures for performing tests and/or calibrations
34 | 4.1 Organization | 4.1.5 General Requirements | Duties - Preventive Actions | Requirement | V1M2 4.1.5 a) | have managerial and technical personnel who have the authority and resources needed to initiate actions to prevent or minimize such departures (see also 5.2);
35 | 4.1 Organization | 4.1.5 General Requirements | Undue Pressures | Procedure | V1M2 4.1.5 b) | have arrangements to ensure that its management and personnel are free from any undue internal and external commercial, financial and other pressures and influences that may adversely affect the quality of their work;
36 | 4.1 Organization | 4.1.5 General Requirements | Confidentiality | Procedure | V1M2 4.1.5 c) | have policies and procedures to ensure the protection of its customers' confidential information and proprietary rights, including procedures for protecting the electronic storage and transmission of results;
37 | 4.1 Organization | 4.1.5 General Requirements | Confidence | Procedure | V1M2 4.1.5 d) | have policies and procedures to avoid involvement in any activities that would diminish confidence in its competence, impartiality, judgment or operational integrity;
38 | 4.1 Organization | 4.1.5 General Requirements | Organization Structure |  | V1M2 4.1.5 e) | define the organization and management structure of the laboratory, its place in any parent organization, and the relationships between quality management, technical operations and support services; (V1M2 4.2.8.4 e)
39 | 4.1 Organization | 4.1.5 General Requirements | Organization Structure |  | V1M2 4.1.5 f) | specify the responsibility, authority and interrelationships of all personnel who manage, perform or verify work affecting the quality of the tests and/or calibrations; (V1M2 4.2.8.4 e and 5.2.5)
40 | 4.1 Organization | 4.1.5 General Requirements | Organization Structure |  | V1M2 4.1.5 g) | provide adequate supervision of testing and calibration staff, including trainees, by persons familiar with methods and procedures, purpose of each test and/or calibration, and with the assessment of the test or calibration results; (see V1M2 4.2.8.4 e and 5.2.1)
41 | 4.1 Organization | 4.1.5 General Requirements | Technical Manager | Requirement | V1M2 4.1.5 h) | have technical management which has overall responsibility for the technical operations and the provision of the resources needed to ensure the required quality of laboratory operations;
42 |  |  |  |  | V1M2 4.1.5 i) | Note: This is different from the actual qualifications of a Quality Manager. Can be determined from the organizational chart, the description of duties, and any documents relating to the selection qualifications and criteria for the position.
43 | 4.1 Organization | 4.1.5 General Requirements | Quality Manager | Requirement | V1M2 4.1.5 i) | appoint a member of staff as quality manager (however named) who, irrespective of other duties and responsibilities, shall have defined responsibility and authority for ensuring that the management system related to quality is implemented and followed at all ; the quality manager shall have direct access to the highest level of management at which decisions are made on laboratory policy or resources;
44 | 4.1 Organization | 4.1.5 General Requirements | Temporary Appointments | Requirement | V1M2 4.1.5 j) | appoint deputies for key managerial personnel (see Note);
45 | 4.1 Organization | 4.1.5 General Requirements | Communication | Requirement | V1M2 4.1.5 k) | ensure that its personnel are aware of the relevance and importance of their activities and how they contribute to the achievement of the objectives of the management system.
46 | 4.1 Organization | 4.1.5 General Requirements |  | Note | V1M2 4.1.5 k) N | NOTE: Individuals may have more than one function and it may be impractical to appoint deputies for every function.
47 | 4.1 Organization | 4.1.6 Communication | Communication | Requirement | V1M2 4.1.6 | Top management shall ensure that appropriate communication processes are established within the laboratory and that communication takes place regarding the effectiveness of the management system.
48 | 4.1 Organization | 4.1.7 Additional General Requirements |  |  | V1M2 4.1.7 | Additional Requirements for Laboratories
49 | 4.1 Organization | 4.1.7.1 Quality Manager | Quality Manager |  | V1M2 4.1.7.1 | The laboratory's quality manager and/or his/her designee(s) shall:
50 |  | 4.1.7.1 Quality Manager | Quality Manager |  |  | See also 4.1.5 i), 4.2.8.2 and 4.14.1
51 | 4.1 Organization | 4.1.7.1 Quality Manager | Responsibility | Requirement | V1M2 4.1.7.1 a) | serve as the focal point for QA/QC and be responsible for the oversight and/or review of quality control data;
52 | 4.1 Organization | 4.1.7.1 Quality Manager | Responsibility | Requirement | V1M2 4.1.7.1 b) | have functions independent from laboratory operations for which they have quality assurance oversight;
53 | 4.1 Organization | 4.1.7.1 Quality Manager | Responsibility | Requirement | V1M2 4.1.7.1 c) | be able to evaluate data objectively and perform assessments without outside (e.g., managerial) influence;
54 | 4.1 Organization | 4.1.7.1 Quality Manager | Experience | Requirement | V1M2 4.1.7.1 d) | have documented training and/or experience in QA/QC procedures and the laboratory’s quality system;
55 | 4.1 Organization | 4.1.7.1 Quality Manager | Familiarity | Requirement | V1M2 4.1.7.1 e) | have a general knowledge of the analytical methods for which data review is performed;
56 | 4.1 Organization | 4.1.7.1 Quality Manager | Responsibility | Requirement | V1M2 4.1.7.1 f) | arrange for or conduct internal audits as per Section 4.14 annually;
57 | 4.1 Organization | 4.1.7.1 Quality Manager | Responsibility | Requirement | V1M2 4.1.7.1 g) | notify laboratory management of deficiencies in the quality system; and
58 | 4.1 Organization | 4.1.7.1 Quality Manager | Responsibility | Requirement | V1M2 4.1.7.1 h) | monitor corrective actions.
59 | 4.1 Organization | 4.1.7.1 Quality Manager | Quality Manager | Note | V1M2 4.1.7.1 h) N | NOTE: Where staffing is limited, the quality manager may also be the technical manager.
60 | 4.1 Organization | 4.1.7.2 Technical Manager | Technical Manager |  | V1M2 4.1.7.2 | The laboratory's technical manager(s), however named, and/or his/her designee(s) shall:
61 | 4.1 Organization | 4.1.7.2 Technical Manager | Duties | Requirement | V1M2 4.1.7.2 a) | be a member of the staff of an environmental laboratory who exercises actual day-to-day supervision of laboratory operations for the appropriate fields of accreditation and reporting of results;
62 | 4.1 Organization | 4.1.7.2 Technical Manager | Experience | Requirement | V1M2 4.1.7.2 b) | be experienced in the fields of accreditation for which the laboratory is seeking accreditation;
63 | 4.1 Organization | 4.1.7.2 Technical Manager | Duties |  | V1M2 4.1.7.2 c) | The laboratory's technical manager(s), however named, and/or his/her designee(s) shall have duties that include:
64 | 4.1 Organization | 4.1.7.2 Technical Manager | Monitoring Quality | Requirement | V1M2 4.1.7.2 c) i | • monitoring standards of performance in quality control and quality assurance, and
65 | 4.1 Organization | 4.1.7.2 Technical Manager | Monitoring Data Validity | Requirement | V1M2 4.1.7.2 c) ii | • monitoring the validity of the analyses performed and data generated in the laboratory to assure reliable data.
66 | 4.1 Organization | 4.1.7.2 Technical Manager | Restriction to one lab | Requirement | V1M2 4.1.7.2 d) | not be the technical manager(s) of more than one accredited environmental laboratory without authorization from the primary Accreditation Body. Circumstances to be considered in the decision to grant such authorization shall include:
67 | 4.1 Organization | 4.1.7.2 Technical Manager | Restriction to one lab | AB Considerations | V1M2 4.1.7.2 d) i | • the extent to which operating hours of the laboratories to be directed overlap,
68 | 4.1 Organization | 4.1.7.2 Technical Manager | Restriction to one lab | AB Considerations | V1M2 4.1.7.2 d) ii | • adequacy of supervision in each laboratory, and
69 | 4.1 Organization | 4.1.7.2 Technical Manager | Restriction to one lab | AB Considerations | V1M2 4.1.7.2 d) iii | • the availability of environmental laboratory services in the area served.
70 | 4.1 Organization | 4.1.7.2 Technical Manager | Temporary Technical Manager | Requirement | V1M2 4.1.7.2 e) | if absent for a period of time exceeding fifteen (15) consecutive calendar days shall designate another full-time staff member meeting the qualifications of the technical manager(s) to temporarily perform this function. If this absence exceeds thirty-five (35) consecutive calendar days, the primary accreditation body shall be notified in writing; and
71 | 4.1 Organization | 4.1.7.2 Technical Manager | Temporary Technical Manager | Requirement | V1M2 4.1.7.2 f) | meet qualification requirements as specified in Section 5.2.6.1.
72 | 4.2 Management |  |  |  | V1M2 4.2 | Management (ISO/IEC 17025:2005(E), Clause 4.2)
73 | 4.2 Management | 4.2.1 Management System |  |  | V1M2 4.2.1 | The laboratory shall establish, implement and maintain a management system appropriate to the scope of its activities.
74 | 4.2 Management | 4.2.1 Management System | Policies and Procedure | Documentation | V1M2 4.2.1 | The laboratory shall document its policies, systems, programmes, procedures and instructions to the extent necessary to assure the quality of the test and/or calibration results.
75 | 4.2 Management | 4.2.1 Management System | Communication & Implementation | Requirement | V1M2 4.2.1 | The system's documentation shall be communicated to, understood by, available to, and implemented by the appropriate personnel. (see also V1M2 4.1.6)
76 |  |  |  |  | V1M2 4.2.2 | (Quality Policy Elements to be included in the Quality Manual - see V1M24.2.8.3 h)
77 | 4.2 Management | 4.2.2 Quality Management System | Quality Policy | Requirement | V1M2 4.2.2 | The laboratory's management system policies related to quality, including a quality policy statement, shall be defined in a quality manual (however named). The overall objectives shall be established, and
78 | 4.2 Management | 4.2.2 Quality Management System | Quality Policy | Requirement | V1M2 4.2.2 | shall be reviewed during management review. The quality policy statement shall be issued under the authority of top management. It shall include at least the following:
79 | 4.2 Management | 4.2.2 Quality Management System | Quality Policy | Requirement | V1M2 4.2.2 a) | the laboratory management's commitment to good professional practice and to the quality of its testing and calibration in servicing its customers;
80 | 4.2 Management | 4.2.2 Quality Management System | Quality Policy | Requirement | V1M2 4.2.2 b) | the management's statement of the laboratory's standard of service;
81 | 4.2 Management | 4.2.2 Quality Management System | Quality Policy | Requirement | V1M2 4.2.2 c) | the purpose of the management system related to quality;
82 | 4.2 Management | 4.2.2 Quality Management System | Quality Manual | Requirement | V1M2 4.2.2 d) | The quality policy statement shall include a requirement that all personnel concerned with testing and calibration activities within the laboratory familiarize themselves with the quality documentation and implement the policies and procedures in their work;
83 | 4.2 Management | 4.2.2 Quality Management System | Quality Manual | Requirement | V1M2 4.2.2 e) | The quality policy statement shall include the laboratory management's commitment to comply with this International Standard and to continually improve the effectiveness of the management system.
84 | 4.2 Management | 4.2.2 Quality Management System | Quality Manual | Note | V1M2 4.2.2 e) N | NOTE: The quality policy statement should be concise and may include the requirement that tests and/or calibrations shall always be carried out in accordance with stated methods and customers' requirements. When the test and/or calibration laboratory is part of a larger organization, some quality policy elements may be in other documents.
85 | 4.2 Management | 4.2.3 Commitment | Management System | Requirement | V1M2 4.2.3 | Top management shall provide evidence of commitment to the development and implementation of the management system and to continually improving its effectiveness.
86 | 4.2 Management | 4.2.4 Communication | Customer Satisfaction | Requirement | V1M2 4.2.4 | Top management shall communicate to the organization the importance of meeting customer requirements as well as statutory and regulatory requirements.
87 | 4.2 Management | 4.2.5 Quality Manual | Procedures | Requirement | V1M2 4.2.5 | The quality manual shall include or make reference to the supporting procedures including technical procedures. It shall outline the structure of the documentation used in the management system.
88 | 4.2 Management | 4.2.6 Quality Manual | Roles & Responsibilities | Requirement | V1M2 4.2.6 | The roles and responsibilities of technical management and the quality manager, including their responsibility for ensuring compliance with this International Standard, shall be defined in the quality manual.
89 | 4.2 Management | 4.2.7 Management System Integrity | Changes | Requirement | V1M2 4.2.7 | Top management shall ensure that the integrity of the management system is maintained when changes to the management system are planned and implemented.
90 | 4.2 Management | 4.2.8 Additional Management System Requirements |  |  | V1M2 4.2.8 | Additional Management System Requirements
91 | 4.2 Management | 4.2.8.1 Data Integrity System | Establishment | Requirement | V1M2 4.2.8.1 | The laboratory shall establish and maintain a documented data integrity system. There are four required elements within a data integrity system.
92 | 4.2 Management | 4.2.8.1 Data Integrity System | Data Integrity |  | V1M2 4.2.8.1 | These are,
93 | 4.2 Management | 4.2.8.1 Data Integrity System | Training |  | V1M2 4.2.8.1 | • Data Integrity training.(see 5.2.7)
94 | 4.2 Management | 4.2.8.1 Data Integrity System | Documentation |  | V1M2 4.2.8.1 | • signed data integrity documentation for all laboratory employees,(see 5.2.7)
95 | 4.2 Management | 4.2.8.1 Data Integrity System | Monitoring | Requirement | V1M2 4.2.8.1 | The laboratory shall have within its data integrity system in-depth, periodic monitoring of data integrity,
96 | 4.2 Management | 4.2.8.1 Data Integrity System | Procedure |  | V1M2 4.2.8.1 | • data integrity procedure documentation (see remainder of 4.2.8.1)
97 | 4.2 Management | 4.2.8.1 Data Integrity System | Approval | Requirement | V1M2 4.2.8.1 | The data integrity procedures shall be signed and dated by top management
98 | 4.2 Management | 4.2.8.1 Data Integrity System | Investigation |  | V1M2 4.2.8.1 | The requirements for data integrity investigation are listed in Section 4.16. The requirements for data integrity training and documentation are listed in Section 5.2.7.
99 | 4.2 Management | 4.2.8.1 Data Integrity System | Annual Review | Requirement | V1M2 4.2.8.1 | Management shall annually review data integrity procedures and update as needed.
100 | 4.2 Management | 4.2.8.1 Data Integrity System | Investigation | Procedure | V1M2 4.2.8.1 a) | Laboratory management shall provide a procedure for confidential reporting of data integrity issues in their laboratory. A primary element of the procedure is to assure confidentiality and a receptive environment in which all employees may privately discuss ethical issues or report items of ethical concern.
101 | 4.2 Management | 4.2.8.1 Data Integrity System | Investigation | Procedure | V1M2 4.2.8.1 b) | In instances of ethical concern, the procedure shall include a process whereby laboratory management is to be informed of the need for any further detailed investigation.
102 | 4.2 Management | 4.2.8.2 Quality Manager | Quality Manager | Requirement | V1M2 4.2.8.2 | The quality manager shall be responsible for maintaining the currency of the quality manual.
103 | 4.2 Management | 4.2.8.3 Quality Manual | Required Contents |  | V1M2 4.2.8.3 | The quality manual shall contain:
104 | 4.2 Management | 4.2.8.3 Quality Manual | Title | Requirement | V1M2 4.2.8.3 a) | document title;
105 | 4.2 Management | 4.2.8.3 Quality Manual | Location | Requirement | V1M2 4.2.8.3 b) | laboratory's full name and address;
106 | 4.2 Management | 4.2.8.3 Quality Manual | Responsible Party Location | Requirement | V1M2 4.2.8.3 c) | name, address (if different from above), and telephone number of individual(s) responsible for the laboratory;
107 | 4.2 Management | 4.2.8.3 Quality Manual | Organization Structure | Requirement | V1M2 4.2.8.3 d) | identification of all major organizational units which are to be covered by this quality manual and the effective date of the version;
108 | 4.2 Management | 4.2.8.3 Quality Manual | Signatories | Requirement | V1M2 4.2.8.3 e) | identification of the laboratory's approved signatories;
109 | 4.2 Management | 4.2.8.3 Quality Manual | Signatures | Requirement | V1M2 4.2.8.3 f) | the signed and dated concurrence (with appropriate names and titles), of all responsible parties including the quality manager(s), technical manager(s), and the agent who is in charge of all laboratory activities, such as the laboratory director or laboratory manager;
110 | 4.2 Management | 4.2.8.3 Quality Manual | Objectives | Requirement | V1M2 4.2.8.3 g) | the objectives of the quality system and contain or reference the laboratory’s policies and procedures;
111 |  |  |  |  | V1M2 4.2.8.3 h) | See V1M2 4.2.2, 4.2.2 a), b) and c) for specifics to be included in quality policy
112 | 4.2 Management | 4.2.8.3 Quality Manual | Quality Policy | Requirement | V1M2 4.2.8.3 h) | the laboratory’s official quality policy statement, which shall include quality system objectives and management’s commitment to ethical laboratory practices and to upholding the requirements of this Standard; and
113 | 4.2 Management | 4.2.8.3 Quality Manual | Table of Contents | Requirement | V1M2 4.2.8.3 i) | a table of contents, and applicable lists of references, glossaries and appendices.
114 | 4.2 Management | 4.2.8.4 Quality Manual | Include or Reference |  | V1M2 4.2.8.4 | The quality manual shall contain or reference:
115 | 4.2 Management | 4.2.8.4 Quality Manual | Maintenance, calibration & verification | Procedure | V1M2 4.2.8.4 a) | all maintenance, calibration and verification procedures used by the laboratory in conducting tests;
116 | 4.2 Management | 4.2.8.4 Quality Manual | Equipment | Requirement | V1M2 4.2.8.4 b) | major equipment and reference measurement standards used as well as the facilities and services used by the laboratory in conducting tests;
117 | 4.2 Management | 4.2.8.4 Quality Manual | Verification | Procedure | V1M2 4.2.8.4 c) | verification practices, which may include inter-laboratory comparisons, proficiency testing programs, use of reference materials and internal quality control schemes;
118 | 4.2 Management | 4.2.8.4 Quality Manual | Reporting | Procedure | V1M2 4.2.8.4 d) | procedures for reporting analytical results;
119 |  |  |  |  | V1M2 4.2.8.4 e) | See V1M2 4.1.5 e) , f) and g) for specific requirements to be discussed.
120 | 4.2 Management | 4.2.8.4 Quality Manual | Organization Structure | Requirement | V1M2 4.2.8.4 e) | the organization and management structure of the laboratory, its place in any parent organization, and relevant organizational charts;
121 |  |  |  |  | V1M2 4.2.8.4 f) | See ViM2 4.13 for specifics information to be discussed
122 | 4.2 Management | 4.2.8.4 Quality Manual | Records Retention | Procedure | V1M2 4.2.8.4 f) | procedures to ensure that all records required under this Standard are retained,
123 | 4.2 Management | 4.2.8.4 Quality Manual | Document Control | Procedure | V1M2 4.2.8.4 f) | procedures for control and maintenance of documentation through a document control system that ensures that all standard operating procedures (SOPs), manuals, or documents clearly indicate the time period during which the procedure or document was in force;
124 |  |  |  |  | V1M2 4.2.8.4 g) | See V1M2 4.1.5 e) for additional information
125 | 4.2 Management | 4.2.8.4 Quality Manual | Job descriptions | Requirement | V1M2 4.2.8.4 g) | job descriptions of key staff and reference to the job descriptions of other laboratory staff;
126 | 4.2 Management | 4.2.8.4 Quality Manual | Measurement Traceability | Procedure | V1M2 4.2.8.4 h) | procedures for achieving traceability of measurements;
127 | 4.2 Management | 4.2.8.4 Quality Manual | Method List | Requirement | V1M2 4.2.8.4 i) | a list of all methods under which the laboratory performs its accredited testing;
128 | 4.2 Management | 4.2.8.4 Quality Manual | Review of New Work | Procedure | V1M2 4.2.8.4 j) | procedures for ensuring that the laboratory reviews all new work to ensure that it has the appropriate facilities and resources before commencing such work;
129 | 4.2 Management | 4.2.8.4 Quality Manual | Sample Handling | Procedure | V1M2 4.2.8.4 k) | procedures for handling samples;
130 | 4.2 Management | 4.2.8.4 Quality Manual | Feedback and Corrective Actions | Procedure | V1M2 4.2.8.4 l) | procedures to be followed for feedback and corrective action whenever testing discrepancies are detected, or departures from documented policies and procedures occur;
131 | 4.2 Management | 4.2.8.4 Quality Manual | Departure from Policies and Procedure | Policy | V1M2 4.2.8.4 m) | policy for permitting departures from documented policies and procedures or from standard specifications;
132 | 4.2 Management | 4.2.8.4 Quality Manual | Complaint Process | Procedure | V1M2 4.2.8.4 n) | procedures for dealing with complaints;
133 | 4.2 Management | 4.2.8.4 Quality Manual | Confidentiality | Procedure | V1M2 4.2.8.4 o) | procedures for protecting confidentiality (including national security concerns), and proprietary rights;
134 | 4.2 Management | 4.2.8.4 Quality Manual | Audits and Data Review | Procedure | V1M2 4.2.8.4 p) | procedures for audits and data review;
135 | 4.2 Management | 4.2.8.4 Quality Manual | Adequate Credentials | Procedure | V1M2 4.2.8.4 q) | procedures for establishing that personnel are adequately experienced in the duties they are expected to carry out and are receiving any needed training; and
136 | 4.2 Management | 4.2.8.4 Quality Manual | Electronic Signatures | Policy | V1M2 4.2.8.4 r) | policy addressing the use of unique electronic signatures, where applicable.
137 | 4.2 Management | 4.2.8.5 SOPs | Standard Operating Procedures | Requirement | V1M2 4.2.8.5 | Laboratories shall maintain SOPs that accurately reflect all phases of current laboratory activities, such as assessing data integrity, corrective actions, handling customer complaints, and all methods.
138 | 4.2 Management | 4.2.8.5 SOPs | Standard Operating Procedures | Option | V1M2 4.2.8.5 a) | These documents, for example, may be equipment manuals provided by the manufacturer, or
139 | 4.2 Management | 4.2.8.5 SOPs | Written SOP Criteria | Requirement | V1M2 4.2.8.5 a) | These documents may be internally written documents with adequate detail to allow someone similarly qualified, other than the analyst, to reproduce the procedures used to generate the test result.
140 | 4.2 Management | 4.2.8.5 SOPs | SOP Accessibility | Requirement | V1M2 4.2.8.5 b) | The relevant SOPs shall be readily accessible to all personnel.
141 | 4.2 Management | 4.2.8.5 SOPs | Identifying Information | Requirement | V1M2 4.2.8.5 c) | Each SOP shall clearly indicate the effective date of the document, the revision number, and the signature(s) of the approving authority.
142 | 4.2 Management | 4.2.8.5 SOPs | Changes to external documents | Requirement | V1M2 4.2.8.5 d) | Documents that contain sufficient information to perform the tests, do not need to be supplemented or rewritten as internal procedures if the documents are written in a way that they can be used as written. Any changes, including the use of a selected option, shall be documented and included in the laboratory’s method records.
143 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods | Requirement | V1M2 4.2.8.5 e) | The laboratory shall have and maintain an SOP for each accredited analyte or method.
144 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods |  | V1M2 4.2.8.5 f) | The SOP may be a copy of a published or referenced method or may be written by the laboratory. In cases where modifications to the published method have been made by the laboratory or where the referenced method is ambiguous or provides insufficient detail, these changes or clarifications shall be clearly described. (4.2.8.5 d))
145 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods |  | V1M2 4.2.8.5 f) | Each method shall include or reference the following topics where applicable:
146 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) ii) | applicable matrix or matrices;
147 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) iii) | limits of detection and quantitation ;
148 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) iv) | scope and application, including parameters to be analyzed;
149 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) v) | summary of the method;
150 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) vi) | definitions;
151 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) vii) | interferences;
152 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) viii) | safety;
153 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) ix) | equipment and supplies;
154 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) x) | reagents and standards;
155 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xi) | sample collection, preservation, shipment and storage;
156 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xii) | quality control;
157 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xiii) | calibration and standardization;
158 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xiv) | procedure;
159 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xv) | data analysis and calculations;
160 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xvi) | method performance;
161 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xvii) | pollution prevention;
162 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xviii) | data assessment and acceptance criteria for quality control measures;
163 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xix) | corrective actions for out-of-control data;
164 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xx) | contingencies for handling out-of-control or unacceptable data;
165 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xxi) | waste management;
166 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xxii) | references; and
167 | 4.2 Management | 4.2.8.5 SOPs | Analytical Methods - Contents | Requirement | V1M2 4.2.8.5 f) xxiii) | any tables, diagrams, flowcharts and validation data.
168 | 4.3 Document Control | Document Control |  |  | V1M2 4.3 | Document Control (ISO/IEC 17025:2005(E), Clause 4.3)
169 | 4.3 Document Control | Document Control |  |  | V1M2 4.3.1 | General
170 | 4.3 Document Control | 4.3.1 General | General | Procedure | V1M2 4.3.1 | The laboratory shall establish and maintain procedures to control all documents that form part of its management system (internally generated or from external sources), such as regulations, standards, other normative documents, test and/or calibration methods, as well as drawings, software, specifications, instructions and manuals.
171 | 4.3 Document Control | 4.3.1 General |  |  | V1M2 4.3.1 N1 | NOTE 1: In this context "document" could be policy statements, procedures, specifications, calibration tables, charts, text books, posters, notices, memoranda, software, drawings, plans, etc. These may be on various media, whether hard copy or electronic, and they may be digital, analog, photographic or written.
172 | 4.3 Document Control | 4.3.1 General |  |  | V1M2 4.3.1 N2 | NOTE 2: The control of data related to testing and calibration is covered in 5.4.7. The control of records is covered in 4.13.
173 | 4.3 Document Control | 4.3.2 Document Approval and Use |  |  | V1M2 4.3.2 | Document Approval and Issue
174 | 4.3 Document Control | 4.3.2 Document Approval and Use | Review and Approval | Requirement | V1M2 4.3.2.1 | All documents issued to personnel in the laboratory as part of the management system shall be reviewed and approved for use by authorized personnel prior to issue.
175 | 4.3 Document Control | 4.3.2 Document Approval and Use | Master List | Requirement | V1M2 4.3.2.1 | A master list or an equivalent document control procedure identifying the current revision status and distribution of documents in the management system shall be established and shall be readily available to preclude the use of invalid and/or obsolete documents.
176 | 4.3 Document Control | 4.3.2 Document Approval and Use | Procedure | Procedure | V1M2 4.3.2.2 | The document procedure(s) adopted shall ensure that:
177 | 4.3 Document Control | 4.3.2 Document Approval and Use | Revisions | Requirement | V1M2 4.3.2.2 a) | authorized editions of appropriate documents are available at all locations where operations essential to the effective functioning of the laboratory are performed;
178 | 4.3 Document Control | 4.3.2 Document Approval and Use | Periodic Review | Requirement | V1M2 4.3.2.2 b) | documents are periodically reviewed and, where necessary, revised to ensure continuing suitability and compliance with applicable requirements;
179 | 4.3 Document Control | 4.3.2 Document Approval and Use | Obsolete Documents | Requirement | V1M2 4.3.2.2 c) | invalid or obsolete documents are promptly removed from all points of issue or use, or otherwise assured against unintended use;
180 | 4.3 Document Control | 4.3.2 Document Approval and Use | Obsolete Documents | Requirement | V1M2 4.3.2.2 d) | obsolete documents retained for either legal or knowledge preservation purposes are suitably marked.
181 | 4.3 Document Control | 4.3.2 Document Approval and Use | Identification | Requirement | V1M2 4.3.2.3 | Management system documents generated by the laboratory shall be uniquely identified.
182 | 4.3 Document Control | 4.3.2 Document Approval and Use | Identification | Requirement | V1M2 4.3.2.3 | Management system document identification shall include the date of issue and/or revision identification, page numbering, the total number of pages or a mark to signify the end of the document, and the issuing authority(ies).
183 | 4.3 Document Control | 4.3.3 Document Changes |  |  | V1M2 4.3.3 | Document Changes
184 | 4.3 Document Control | 4.3.3 Document Changes | Approval | Requirement | V1M2 4.3.3.1 | Changes to documents shall be reviewed and approved by the same function that performed the original review unless specifically designated otherwise.
185 | 4.3 Document Control | 4.3.3 Document Changes | Approval | Requirement | V1M2 4.3.3.1 | The designated personnel shall have access to pertinent background information upon which to base their review and approval.
186 | 4.3 Document Control | 4.3.3 Document Changes | Change Identification | Requirement | V1M2 4.3.3.2 | Where practicable, the altered or new text shall be identified in the document or the appropriate attachments.
187 | 4.3 Document Control | 4.3.3 Document Changes | Amendments | Procedure | V1M2 4.3.3.3 | If the laboratory's document control system allows for the amendment of documents by hand pending the re-issue of the documents, the procedures and authorities for such amendments shall be defined.
188 | 4.3 Document Control | 4.3.3 Document Changes | Amendments | Requirement | V1M2 4.3.3.3 | Amendments shall be clearly marked, initialled and dated.
189 | 4.3 Document Control | 4.3.3 Document Changes | Amendments | Requirement | V1M2 4.3.3.3 | A revised document shall be formally re-issued as soon as practicable.
190 | 4.3 Document Control | 4.3.3 Document Changes | Electronic Records | Procedure | V1M2 4.3.3.4 | Procedures shall be established to describe how changes in documents maintained in computerized systems are made and controlled.
191 | 4.4 Review of Requests, Tenders and Contracts |  |  |  | V1M2 4.4 | Review of Requests, Tenders and Contracts (ISO/IEC 17025:2005(E), Clause 4.4)
192 | 4.4 Review of Requests, Tenders and Contracts | 4.4.1 Procedures for Review |  |  | V1M2 4.4.1 | The laboratory shall establish and maintain procedures for the review of requests, tenders and contracts. (4.4.1 a) through c)
193 | 4.4 Review of Requests, Tenders and Contracts | 4.4.1 Procedures for Review | Review Objectives | Requirement | V1M2 4.4.1 a) | The policies and procedures for the review of requests, tenders and contracts leading to a contract for testing and/or calibration shall ensure that: the requirements, including the methods to be used, are adequately defined, documented and understood (see 5.4.2);
194 | 4.4 Review of Requests, Tenders and Contracts | 4.4.1 Procedures for Review | Capacity and Resources | Requirement | V1M2 4.4.1 b) | The policies and procedures for the review of requests, tenders and contracts leading to a contract for testing and/or calibration shall ensure that: the laboratory has the capability and resources to meet the requirements;
195 | 4.4 Review of Requests, Tenders and Contracts | 4.4.1 Procedures for Review | Appropriate Methods | Requirement | V1M2 4.4.1 c) | The policies and procedures for he review of requests, tenders and contracts leading to a contract for testing and/or calibration shall ensure that: the appropriate test and/or calibration method is selected and is capable of meeting the customers' requirements (see 5.4.2).
196 | 4.4 Review of Requests, Tenders and Contracts | 4.4.1 Procedures for Review | Resolving Differences | Requirement | V1M2 4.4.1 c) | Any differences between the request or tender and the contract shall be resolved before any work commences. Each contract shall be acceptable both to the laboratory and the customer.
197 | 4.4 Review of Requests, Tenders and Contracts | 4.4.1 Procedures for Review |  |  | V1M2 4.4.1 c) N1 | NOTE 1: The request, tender and contract review should be conducted in a practical and efficient manner, and the effect of financial, legal and time schedule aspects should be taken into account. For internal customers, reviews of requests, tenders and contracts can be performed in a simplified way.
198 | 4.4 Review of Requests, Tenders and Contracts | 4.4.1 Procedures for Review |  |  | V1M2 4.4.1 c) N2 | NOTE 2: The review of capability should establish that the laboratory possesses the necessary physical, personnel and information resources, and that the laboratory's personnel have the skills and expertise necessary for the performance of the tests and/or calibrations in question. The review may also encompass results of earlier participation in interlaboratory comparisons or proficiency testing and/or the running of trial test or calibration programmes using samples or items of known value in order to determine uncertainties of measurement, limits of detection, confidence limits, etc.
199 | 4.4 Review of Requests, Tenders and Contracts | 4.4.1 Procedures for Review |  |  | V1M2 4.4.1 c) N3 | NOTE 3: A contract may be any written or oral agreement to provide a customer with testing and/or calibration services.
200 | 4.4 Review of Requests, Tenders and Contracts | 4.4.2 Records Review | Review | Requirement | V1M2 4.4.2 | Records of the review of requests, tenders and contracts, including any significant changes, shall be maintained.
201 | 4.4 Review of Requests, Tenders and Contracts | 4.4.2 Records Review | Customer Communication | Record | V1M2 4.4.2 | Records shall also be maintained of pertinent discussions with a customer relating to the customer's requirements or the results of the work during the period of execution of the contract.
202 | 4.4 Review of Requests, Tenders and Contracts | Review of Requests, Tenders and Contracts |  |  | V1M2 4.4.2 N | NOTE: For review of routine and other simple tasks, the date and the identification (e.g. the initials) of the person in the laboratory responsible for carrying out the contracted work are considered adequate. For repetitive routine tasks, the review need be made only at the initial enquiry stage or on granting of the contract for on-going routine work performed under a general agreement with the customer, provided that the customer's requirements remain unchanged. For new, complex or advanced testing and/or calibration tasks, a more comprehensive record should be maintained.
203 | 4.4 Review of Requests, Tenders and Contracts | 4.4.3 Subcontracted Laboratory | Review | Requirement | V1M2 4.4.3 | The review of requests, tenders and contracts shall also cover any work that is subcontracted by the laboratory.
204 | 4.4 Review of Requests, Tenders and Contracts | 4.4.4 Customer Communication | Deviations | Requirement | V1M2 4.4.4 | The customer shall be informed of any deviation from the contract.
205 | 4.4 Review of Requests, Tenders and Contracts | 4.4.5 Customer Communication | Amendments | Requirement | V1M2 4.4.5 | If a contract needs to be amended after work has commenced, the same contract review process shall be repeated and any amendments shall be communicated to all affected personnel.
206 | 4.5 Subcontracting tests |  |  |  | V1M2 4.5 | Subcontracting of Environmental Tests (ISO/IEC 17025:2005(E), Clause 4.5)
207 | 4.5 Subcontracting tests | Subcontracting Tests |  |  | V1M2 4.5.1 | When a laboratory subcontracts work, whether because of unforeseen reasons (e.g. workload, need for further expertise or temporary incapacity) or on a continuing basis (e.g. through permanent subcontracting, agency or franchising arrangements), this work shall be placed with a competent subcontractor. A competent subcontractor is one that, for example, complies with this International Standard for the work in question.
208 | 4.5 Subcontracting tests | Customer Obligation | Communication | Requirement | V1M2 4.5.2 | The laboratory shall advise the customer of the arrangement in writing and, when appropriate, gain the approval of the customer, preferably in writing.
209 | 4.5 Subcontracting tests | Customer Obligation | Responsibility | Requirement | V1M2 4.5.3 | The laboratory is responsible to the customer for the subcontractor's work, except in the case where the customer or a regulatory authority specifies which subcontractor is to be used.
210 | 4.5 Subcontracting tests | Subcontractor use and compliance | Registry | Record | V1M2 4.5.4 | The laboratory shall maintain a register of all subcontractors that it uses for tests and/or calibrations and a record of the evidence of compliance with this International Standard for the work in question.
211 | 4.5 Subcontracting tests | Subcontracted Laboratory | Credentials | Requirement | V1M2 4.5.5 | When a laboratory subcontracts work, this work shall be placed with a laboratory accredited to this Standard for the tests to be performed or with a laboratory that meets applicable statutory and regulatory requirements for performing the tests and submitting the results of tests performed.
212 | 4.5 Subcontracting tests | Subcontracted Laboratory | Identification | Requirement | V1M2 4.5.5 | The laboratory performing the subcontracted work shall be indicated in the final report.
213 | 4.5 Subcontracting tests | Subcontracted Laboratory | Final Report | Record | V1M2 4.5.5 | The laboratory shall make a copy of the subcontractor’s report available to the client when requested.
214 | 4.6 Purchasing Services & Supplies |  |  |  | V1M2 4.6 | Purchasing Services and Supplies (ISO/IEC 17025:2005(E), Clause 4.6)
215 | 4.6 Purchasing Services & Supplies | Procedure Requirement | Selection and Purchasing | Procedure | V1M2 4.6.1 | The laboratory shall have a policy and procedure(s) for the selection and purchasing of services and supplies it uses that affect the quality of the tests and/or calibrations.
216 | 4.6 Purchasing Services & Supplies | Procedure Requirement | Reception and Storage | Procedure | V1M2 4.6.1 | Procedures shall exist for the purchase, reception and storage of reagents and laboratory consumable materials relevant for the tests and calibrations.
217 | 4.6 Purchasing Services & Supplies | Assessing Purchases | Verification | Requirement | V1M2 4.6.2 | The laboratory shall ensure that purchased supplies and reagents and consumable materials that affect the quality of tests and/or calibrations are not used until they have been inspected or otherwise verified as complying with standard specifications or requirements defined in the methods for the tests and/or calibrations concerned.
218 | 4.6 Purchasing Services & Supplies | Assessing Purchases | Specifications | Requirement | V1M2 4.6.2 | These services and supplies used shall comply with specified requirements.
219 | 4.6 Purchasing Services & Supplies | Assessing Purchases | Specifications | Record | V1M2 4.6.2 | Records of actions taken to check compliance of services and supplies shall be maintained.
220 | 4.6 Purchasing Services & Supplies | Assessing Purchases | Purchasing Documents | Record | V1M2 4.6.3 | Purchasing documents for items affecting the quality of laboratory output shall contain data describing the services and supplies ordered.
221 | 4.6 Purchasing Services & Supplies | Assessing Purchases | Review | Requirement | V1M2 4.6.3 | These purchasing documents shall be reviewed and approved for technical content prior to release.
222 | 4.6 Purchasing Services & Supplies |  |  |  | V1M2 4.6.3 N | NOTE: The description may include type, class, grade, precise identification, specifications, drawings, inspection instructions, other technical data including approval of test results, the quality required and the management system standard under which they were made.
223 | 4.6 Purchasing Services & Supplies | Suppliers | Evaluation | Requirement | V1M2 4.6.4 | The laboratory shall evaluate suppliers of critical consumables, supplies and services which affect the quality of testing and calibration,
224 | 4.6 Purchasing Services & Supplies | Suppliers | Evaluation | Record | V1M2 4.6.4 | The laboratory shall maintain records of the evaluation suppliers of critical consumables, supplies and services and list those approved.
225 | 4.7 Service to Clients |  |  |  | V1M2 4.7 | Service to the Client (ISO/IEC 17025:2005(E), Clause 4.7)
226 | 4.7 Service to Clients | General | Cooperation | Requirement | V1M2 4.7.1 | The laboratory shall be willing to cooperate with customers or their representatives in clarifying the customer's request and in monitoring the laboratory's performance in relation to the work performed, provided that the laboratory ensures confidentiality to other customers.
227 | 4.7 Service to Clients |  |  |  | V1M2 4.7.1 N1 | NOTE 1: Such cooperation may include:
228 | 4.7 Service to Clients |  |  |  | V1M2 4.7.1 N1 a) | providing the customer or the customer's representative reasonable access to relevant areas of the laboratory for the witnessing of tests and/or calibrations performed for the customer;
229 | 4.7 Service to Clients |  |  |  | V1M2 4.7.1 N1 b) | preparation, packaging, and dispatch of test and/or calibration items needed by the customer for verification purposes.
230 | 4.7 Service to Clients |  |  |  | V1M2 4.7.1 N2 | NOTE 2: Customers value the maintenance of good communication, advice and guidance in technical matters, and opinions and interpretations based on results. Communication with the customer, especially in large assignments, should be maintained throughout the work. The laboratory should inform the customer of any delays or major deviations in the performance of the tests and/or calibrations.
231 | 4.7 Service to Clients | Customer Communication | Feedback | Requirement | V1M2 4.7.2 | The laboratory shall seek feedback, both positive and negative, from its customers.
232 | 4.7 Service to Clients | Customer Communication | Using Feedback | Requirement | V1M2 4.7.2 | The feedback shall be used and analysed to improve the management system, testing and calibration activities and customer service.
233 | 4.7 Service to Clients |  |  |  | V1M2 4.7.2 N | NOTE: Examples of the types of feedback include customer satisfaction surveys and review of test or calibration reports with customers.
234 | 4.8 Complaints | Complaints |  |  | V1M2 4.8 | Complaints (ISO/IEC 17025:2005(E), Clause 4.8)
235 | 4.8 Complaints | General | Complaint Resolution | Procedure | V1M2 4.8 | The laboratory shall have a policy and procedure for the resolution of complaints received from customers or other parties.
236 | 4.8 Complaints | General | Complaint Resolution | Record | V1M2 4.8 | Records shall be maintained of all complaints and of the investigations and corrective actions taken by the laboratory (see also 4.11).
237 | 4.9 Control of Noncoforming Work |  |  |  | V1M2 4.9 | Control of Nonconforming Environmental Testing Work (ISO/IEC 17025:2005(E), Clause 4.9)
238 | 4.9 Control of Noncoforming Work | Nonconforming Work | Implementation | Procedure | V1M2 4.9.1 | The laboratory shall have a policy and procedures that shall be implemented when any aspect of its testing and/or calibration work, or the results of this work, do not conform to its own procedures or the agreed requirements of the customer. The policy and procedures shall ensure that:
239 | 4.9 Control of Noncoforming Work | Nonconforming Work | Implementation | Procedure | V1M2 4.9.1 | The policy and procedures shall ensure that:
240 | 4.9 Control of Noncoforming Work | Implementing Procedures | Responsibility | Procedure | V1M2 4.9.1 a) | the responsibilities and authorities for the management of nonconforming work are designated and actions (including halting of work and withholding of test reports and calibration certificates, as necessary) are defined and taken when nonconforming work is identified;
241 | 4.9 Control of Noncoforming Work | Implementing Procedures | Evaluation | Requirement | V1M2 4.9.1 b) | an evaluation of the significance of the nonconforming work is made;
242 | 4.9 Control of Noncoforming Work | Implementing Procedures | Corrective Action | Requirement | V1M2 4.9.1 c) | correction is taken immediately, together with any decision about the acceptability of the nonconforming work;
243 | 4.9 Control of Noncoforming Work | Implementing Procedures | Notification | Requirement | V1M2 4.9.1 d) | where necessary, the customer is notified and work is recalled;
244 | 4.9 Control of Noncoforming Work | Implementing Procedures | Work Resumption | Requirement | V1M2 4.9.1 e) | the responsibility for authorizing the resumption of work is defined.
245 | 4.9 Control of Noncoforming Work | Nonconforming Work |  |  | V1M2 4.9.1 e) N | NOTE: Identification of nonconforming work or problems with the management system or with testing and/or calibration activities can occur at various places within the management system and technical operations. Examples are customer complaints, quality control, instrument calibration, checking of consumable materials, staff observations or supervision, test report and calibration certificate checking, management reviews and internal or external audits.
246 | 4.9 Control of Noncoforming Work | Evaluation Results | Corrective Action | Requirement | V1M2 4.9.2 | Where the evaluation indicates that the nonconforming work could recur or that there is doubt about the compliance of the laboratory's operations with its own policies and procedures, the corrective action procedures given in 4.11 shall be promptly followed.
247 | 4.10 Improvement | Improvement |  |  | V1M2 4.1 | Improvement (ISO/IEC 17025:2005(E), Clause 4.10)
248 | 4.10 Improvement | Continuous Improvement | Tools | Requirement | V1M2 4.1 | The laboratory shall continually improve the effectiveness of its management system through the use of the quality policy, quality objectives, audit results, analysis of data, corrective and preventive actions and management review.
249 | 4.11 Corrective Actions | Corrective Actions |  |  | V1M2 4.11 | Corrective Action (ISO/IEC 17025:2005(E), Clause 4.11)
250 | 4.11 Corrective Actions | General |  |  | V1M2 4.11.1 | General
251 | 4.11 Corrective Actions | General | Implementing Action | Procedure | V1M2 4.11.1 | The laboratory shall establish a policy and a procedure for implementing corrective action when nonconforming work or departures from the policies and procedures in the management system or technical operations have been identified.
252 | 4.11 Corrective Actions | General | Appropriate Authorities | Requirement | V1M2 4.11.1 | The laboratory shall designate appropriate authorities for implementing corrective action when nonconforming work or departures from the policies and procedures in the management system or technical operations have been identified.
253 | 4.11 Corrective Actions |  |  |  | V1M2 4.11.1 N | NOTE: A problem with the management system or with the technical operations of the laboratory may be identified through a variety of activities, such as control of nonconforming work, internal or external audits, management reviews, feedback from customers and from staff observations.
254 | 4.11 Corrective Actions | Cause Analysis | Cause Analysis |  | V1M2 4.11.2 | Cause Analysis
255 | 4.11 Corrective Actions | Cause Analysis | Investigation | Procedure | V1M2 4.11.2 | The procedure for corrective action shall start with an investigation to determine the root cause(s) of the problem.
256 | 4.11 Corrective Actions | Cause Analysis |  |  | V1M2 4.11.2 N | NOTE: Cause analysis is the key and sometimes the most difficult part in the corrective action procedure. Often the root cause is not obvious and thus a careful analysis of all potential causes of the problem is required. Potential causes could include customer requirements, the samples, sample specifications, methods and procedures, staff skills and training, consumables, or equipment and its calibration.
257 | 4.11 Corrective Actions | Selection and Implementation |  |  | V1M2 4.11.3 | Selection and Implementation of Corrective Actions
258 | 4.11 Corrective Actions | Selection and Implementation | Identification | Requirement | V1M2 4.11.3 | Where corrective action is needed, the laboratory shall identify potential corrective actions
259 | 4.11 Corrective Actions | Selection and Implementation | Selection | Requirement | V1M2 4.11.3 | It shall select and implement the action(s) most likely to eliminate the problem and to prevent recurrence.
260 | 4.11 Corrective Actions | Selection and Implementation | Appropriateness | Requirement | V1M2 4.11.3 | Corrective actions shall be to a degree appropriate to the magnitude and the risk of the problem.
261 | 4.11 Corrective Actions | Selection and Implementation | Changes | Record | V1M2 4.11.3 | The laboratory shall document and implement any required changes resulting from corrective action investigations.
262 | 4.11 Corrective Actions | Monitoring Corrective Actions | Monitoring |  | V1M2 4.11.4 | Monitoring of Corrective Actions
263 | 4.11 Corrective Actions | Monitoring Corrective Actions | Monitoring | Requirement | V1M2 4.11.4 | The laboratory shall monitor the results to ensure that the corrective actions taken have been effective.
264 | 4.11 Corrective Actions | Additional Audits | Additional Audits |  | V1M2 4.11.5 | Additional Audits
265 | 4.11 Corrective Actions | Additional Audits | Audit Scheduling | Requirement | V1M2 4.11.5 | Where the identification of nonconformities or departures casts doubts on the laboratory's compliance with its own policies and procedures, or on its compliance with this International Standard, the laboratory shall ensure that the appropriate areas of activity are audited in accordance with 4.14 as soon as possible.
266 | 4.11 Corrective Actions | Additional Audits | Additional Audits |  | V1M2 4.11.5 N | NOTE: Such additional audits often follow the implementation of the corrective actions to confirm their effectiveness. An additional audit should be necessary only when a serious issue or risk to the business is identified.
267 | 4.11 Corrective Actions | Documented Procedures | Corrective Action | Procedure | V1M2 4.11.6 | The laboratory shall have documented procedure(s) to address 4.11.1 (see 4.11.1)
268 | 4.11 Corrective Actions | Documented Procedures | Implementation | Procedure | V1M2 4.11.6 | The laboratory shall have documented procedure(s) to address 4.11.3
269 | 4.11 Corrective Actions | Documented Procedures | Monitoring | Procedure | V1M2 4.11.6 | The laboratory shall have documented procedure(s) to address 4.11.4
270 | 4.11 Corrective Actions | Documented Procedures | Scheduling | Procedure | V1M2 4.11.6 | The laboratory shall have documented procedure(s) to address 4.11.5
271 | 4.11 Corrective Actions | Documented Procedures | QC Assessment | Requirement | V1M2 4.11.6 a) | These procedure(s) shall also include: which individual(s) or positions are responsible for assessing each QC data type; and
272 | 4.11 Corrective Actions | Documented Procedures | Initiating CA | Procedure | V1M2 4.11.6 b) | These procedure(s) shall also include: which individual(s) or positions are responsible for initiating and/or recommending corrective actions.
273 | 4.11 Corrective Actions | Cause Analysis | Use | Requirement | V1M2 4.11.7 | Cause analysis described in 4.11.2 applies to failures that indicate a systematic error.
274 | 4.12 Preventive Action | Preventive Actions |  |  | V1M2 4.12 | Preventive Action (ISO/IEC 17025:2005(E), Clause 4.12)
275 | 4.12 Preventive Action | General | Identification | Requirement | V1M2 4.12.1 | Needed improvements and potential sources of nonconformities, either technical or concerning the management system, shall be identified.
276 | 4.12 Preventive Action | General | Implementation | Requirement | V1M2 4.12.1 | When improvement opportunities are identified or if preventive action is required, action plans shall be developed, implemented and monitored to reduce the likelihood of the occurrence of such nonconformities and to take advantage of the opportunities for improvement.
277 | 4.12 Preventive Action | General | Description | Procedure | V1M2 4.12.2 | Procedures for preventive actions shall include the initiation of such actions and the application of controls to ensure that they are effective.
278 | 4.12 Preventive Action | Preventive Actions |  |  | V1M2 4.12.2 N1 | NOTE 1: Preventive action is a pro-active process to identify opportunities for improvement rather than a reaction to the identification of problems or complaints.
279 | 4.12 Preventive Action | Preventive Actions |  |  | V1M2 4.12.2 N2 | NOTE 2: Apart from the review of the operational procedures, the preventive action might involve analysis of data, including trend and risk analyses and proficiency-testing results.
280 | 4.13 Control of Records | Control of Records |  |  | V1M2 4.13 | Control of Records (ISO/IEC 17025:2005(E), Clause 4.13)
281 | 4.13 Control of Records | General |  |  | V1M2 4.13.1 | General
282 | 4.13 Control of Records | General | Scope | Procedure | V1M2 4.13.1.1 | The laboratory shall establish and maintain procedures for identification, collection, indexing, access, filing, storage, maintenance and disposal of quality and technical records.
283 | 4.13 Control of Records | General | Quality Records | Requirement | V1M2 4.13.1.1 | Quality records shall include reports from internal audits and management reviews as well as records of corrective and preventive actions.
284 | 4.13 Control of Records | General | Storage | Requirement | V1M2 4.13.1.2 | All records shall be legible and shall be stored and retained in such a way that they are readily retrievable in facilities that provide a suitable environment to prevent damage or deterioration and to prevent loss.
285 | 4.13 Control of Records | General | Retention Time | Requirement | V1M2 4.13.1.2 | Retention times of records shall be established.
286 | 4.13 Control of Records | General | General |  | V1M2 4.13.1.2 N | NOTE: Records may be in any media, such as hard copy or electronic media.
287 | 4.13 Control of Records | General | Confidentiality | Requirement | V1M2 4.13.1.3 | All records shall be held secure and in confidence.
288 | 4.13 Control of Records | General | Electronic Records | Procedure | V1M2 4.13.1.4 | The laboratory shall have procedures to protect and back-up records stored electronically
289 | 4.13 Control of Records | General | Unauthorized Access | Procedure | V1M2 4.13.1.4 | The laboratory shall have procedures to prevent unauthorized access to or amendment of these records.
290 | 4.13 Control of Records | 4.13.2 Technical Records | Technical |  | V1M2 4.13.2 | Technical Records
291 | 4.13 Control of Records | 4.13.2 Technical Records | Record Identification | Requirement | V1M2 4.13.2.1 | The laboratory shall retain records of original observations, derived data and sufficient information to establish an audit trail, calibration records, staff records and a copy of each test report or calibration certificate issued, for a defined period.
292 | 4.13 Control of Records | 4.13.2 Technical Records | Historical Reconstruction | Requirement | V1M2 4.13.2.1 | The records for each test or calibration shall contain sufficient information to facilitate, if possible, identification of factors affecting the uncertainty and to enable the test or calibration to be repeated under conditions as close as possible to the original.(4.13.3 f)
293 | 4.13 Control of Records | 4.13.2 Technical Records | Personnel Identification | Requirement | V1M2 4.13.2.1 | The records shall include the identity of personnel responsible for the sampling, performance of each test and/or calibration and checking of results.
294 | 4.13 Control of Records | 4.13.2 Technical Records |  |  | V1M2 4.13.2.1 N1 | NOTE 1: In certain fields it may be impossible or impractical to retain records of all original observations.
295 | 4.13 Control of Records | 4.13.2 Technical Records |  |  | V1M2 4.13.2.1 N2 | NOTE 2: Technical records are accumulations of data (see 5.4.7) and information which result from carrying out tests and/or calibrations and which indicate whether specified quality or process parameters are achieved. They may include forms, contracts, work sheets, work books, check sheets, work notes, control graphs, external and internal test reports and calibration certificates, customers' notes, papers and feedback.
296 | 4.13 Control of Records | 4.13.2 Technical Records | Recording Data | Requirement | V1M2 4.13.2.2 | Observations, data and calculations shall be recorded at the time they are made and shall be identifiable to the specific task.
297 | 4.13 Control of Records | 4.13.2 Technical Records | Error Correction | Requirement | V1M2 4.13.2.3 | When mistakes occur in records, each mistake shall be crossed out, not erased, made illegible or deleted, and the correct value entered alongside.
298 | 4.13 Control of Records | 4.13.2 Technical Records | Error Correction | Requirement | V1M2 4.13.2.3 | All such alterations to records shall be signed or initialled by the person making the correction.
299 | 4.13 Control of Records | 4.13.2 Technical Records | Error Correction | Requirement | V1M2 4.13.2.3 | In the case of records stored electronically, equivalent measures shall be taken to avoid loss or change of original data.
300 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements |  |  | V1M2 4.13.3 | Additional Requirements
301 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Requirement | V1M2 4.13.3 a) | The laboratory shall establish a record keeping system that allows the history of the sample and associated data to be readily understood through the documentation.
302 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Requirement | V1M2 4.13.3 a) | This system shall produce unequivocal, accurate records that document all laboratory activities such as laboratory facilities, equipment, analytical methods, and related laboratory activities, such as sample receipt, sample preparation, or data verification, and inter-laboratory transfers of samples and/or extracts.
303 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Records Retention | Requirement | V1M2 4.13.3 b) | The laboratory shall retain all records for a minimum of five years from generation of the last entry in the records.
304 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Availability | Requirement | V1M2 4.13.3 c) | Records shall be available to the accreditation body.
305 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Electronic Records | Requirement | V1M2 4.13.3 d) | Records that are stored only on electronic media shall be supported by the hardware and software necessary for their retrieval.
306 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Archived Records | Requirement | V1M2 4.13.3 e) | Access to archived information shall be documented with an access log.
307 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Requirement | V1M2 4.13.3 f) | All information necessary for the historical reconstruction of data shall be maintained by the laboratory.
308 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Requirement | V1M2 4.13.3 f) | Historical reconstruction information shall include
309 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) i | all raw data, whether hard copy or electronic, for calibrations, samples and quality control measures, including analysts’ worksheets and data output records (chromatograms, strip charts, and other instrument response readout records);
310 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) ii | a written description or reference to the specific test method used, which includes a description of the specific computational steps used to translate parametric observations into a reportable analytical value;
311 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) iii | laboratory sample ID code;
312 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) iv | date of analysis;
313 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) v | time of analysis is required if the holding time is seventy-two hours or less
314 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) v | time of analysis is required when time critical steps are included in the analysis (e.g., extractions and incubations);
315 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) vi | instrumentation identification and instrument operating conditions/parameters (or reference to such data);
316 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) vii | all manual calculations;
317 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) viii | analyst's or operator's initials/signature or electronic identification;
318 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) ix | sample preparation, including cleanup, separation protocols, incubation periods or subculture, ID codes, volumes, weights, instrument printouts, meter readings, calculations, reagents;
319 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) x | test results;
320 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) xi | standard and reagent origin, receipt, preparation, and use;
321 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) xii | calibration criteria, frequency and acceptance criteria;
322 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) xiii | data and statistical calculations, review, confirmation, interpretation, assessment and reporting conventions;
323 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) xiv | quality control protocols and assessment;
324 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) xv | electronic data security, software documentation and verification, software and hardware audits, backups, and records of any changes to automated data entries;
325 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) xvi | method performance criteria including expected quality control requirements;
326 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) xvii | proficiency test results;
327 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) xviii | records of demonstration of capability for each analyst; and
328 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Historical Reconstruction | Record | V1M2 4.13.3 f) xix | a record of names, initials, and signatures for all individuals who are responsible for signing or initialing any laboratory record.
329 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Legible Records | Requirement | V1M2 4.13.3 g | All generated data, except those that are generated by automated data collection systems, shall be recorded legibly in permanent ink.
330 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Error Correction | Requirement | V1M2 4.13.3 g i | i. An individual making corrections to records shall date and initial the correction.
331 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Error Correction | Requirement | V1M2 4.13.3 g ii | ii. Corrections due to reasons other than transcription errors shall specify the reason for the correction.
332 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | Records Transfer | Requirement | V1M2 4.13.3 h | The laboratory shall have a plan to ensure that the records are maintained or transferred according to the clients’ instructions in the event that a laboratory transfers ownership or goes out of business.
333 | 4.13 Control of Records | 4.13.3 Technical Records Additional Requirements | State Requirements | State Requirements | V1M2 4.13.3 h | In addition, appropriate regulatory and state legal requirements concerning laboratory records shall be followed.
334 | 4.14 Internal Audits | Internal Audits |  |  | V1M2 4.14 | Internal Audits (ISO/IEC 17025:2005(E), Clause 4.14)
335 | 4.14 Internal Audits | General | Schedule | Requirement | V1M2 4.14.1 | The laboratory shall periodically, and in accordance with a predetermined schedule and procedure, conduct internal audits of its activities to verify that its operations continue to comply with the requirements of the management system and this International Standard.
336 | 4.14 Internal Audits | General | Scope | Requirement | V1M2 4.14.1 | The internal audit programme shall address all elements of the management system, including the testing and/or calibration activities.
337 | 4.14 Internal Audits | General | Scheduling Responsibility | Requirement | V1M2 4.14.1 | It is the responsibility of the quality manager to plan and organize audits as required by the schedule and requested by management.
338 | 4.14 Internal Audits | General | Auditor Qualifications | Requirement | V1M2 4.14.1 | Such audits shall be carried out by trained and qualified personnel who are, wherever resources permit, independent of the activity to be audited.
339 | 4.14 Internal Audits | Internal Audits |  |  | V1M2 4.14.1 N | NOTE: The cycle for internal auditing should normally be completed in one year.
340 | 4.14 Internal Audits | Audit Results | Need for Corrective Actions | Requirement | V1M2 4.14.2 | When audit findings cast doubt on the effectiveness of the operations or on the correctness or validity of the laboratory's test or calibration results, the laboratory shall take timely corrective action
341 | 4.14 Internal Audits | Audit Results | Customer Notification | Requirement | V1M2 4.14.2 | When audit findings cast doubt on the effectiveness of the operations or on the correctness or validity of the laboratory's test or calibration results, the laboratory notify customers in writing if investigations show that the laboratory results may have been affected.
342 | 4.14 Internal Audits | Audit Description | Record Contents | Record | V1M2 4.14.3 | The area of activity audited, the audit findings and corrective actions that arise from them shall be recorded.
343 | 4.14 Internal Audits | Corrective Actions | Monitoring | Requirement | V1M2 4.14.4 | Follow-up audit activities shall verify and record the implementation and effectiveness of the corrective action taken.
344 | 4.14 Internal Audits | Additional Items |  |  | V1M2 4.14.5 | Additional Items
345 | 4.14 Internal Audits | Additional Items | Customer Notification | Requirement | V1M2 4.14.5 a) | The laboratory shall have a policy that specifies the time frame for notifying a client of events that cast doubt on the validity of the results.
346 | 4.14 Internal Audits | Additional Items | Customer Notification | Requirement | V1M2 4.14.5 b) | The laboratory management shall ensure that these actions are discharged within the agreed time frame.
347 | 4.14 Internal Audits | Additional Items | Schedule | Requirement | V1M2 4.14.5 c) | The Internal audit schedule shall be completed annually,
348 | 4.15 Management Reviews |  |  |  | V1M2 4.15 | Management Reviews (ISO/IEC 17025:2005(E), Clause 4.15)
349 | 4.15 Management Reviews | General | Schedule | Requirement | V1M2 4.15.1 | In accordance with a predetermined schedule and procedure, the laboratory's top management shall periodically conduct a review of the laboratory's management system and testing and/or calibration activities to ensure their continuing suitability and effectiveness
350 | 4.15 Management Reviews | General | Schedule | Requirement | V1M2 4.15.1 | In accordance with a predetermined schedule and procedure, the laboratory's top management shall periodically conduct a review of the laboratory's management system and testing and/or calibration activities to introduce necessary changes or improvements.
351 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The management review shall take account of: the suitability of policies and procedures;
352 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The management review shall take account of: reports from managerial and supervisory personnel;
353 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The management review shall take account of: the outcome of recent internal audits;
354 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The management review shall take account of: corrective and preventive actions;
355 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The management review shall take account of: assessments by external bodies;
356 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The management review shall take account of: the results of interlaboratory comparisons or proficiency tests;
357 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The review shall take account of: changes in the volume and type of the work;
358 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The management review shall take account of: customer feedback;
359 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The management review shall take account of: complaints;
360 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The management review shall take account of: recommendations for improvement;
361 | 4.15 Management Reviews | General | Review Content | Requirement | V1M2 4.15.1 | -The management review shall take account of: other relevant factors, such as quality control activities, resources, and staff training.
362 | 4.15 Management Reviews | Management Reviews |  |  | V1M2 4.15.1 N1 | NOTE 1: A typical period for conducting a management review is once every 12 months.
363 | 4.15 Management Reviews | Management Reviews |  |  | V1M2 4.15.1 N2 | NOTE 2: Results should feed into the laboratory planning system and should include the goals, objectives and action plans for the coming year.
364 | 4.15 Management Reviews | Management Reviews |  |  | V1M2 4.15.1 N3 | NOTE 3: A management review includes consideration of related subjects at regular management meetings.
365 | 4.15 Management Reviews | General | Findings | Record | V1M2 4.15.2 | Findings from management reviews and the actions that arise from them shall be recorded.
366 | 4.15 Management Reviews | General | Action Implementation | Requirement | V1M2 4.15.2 | The management shall ensure that those actions are carried out within an appropriate and agreed timescale.
367 | 4.15 Management Reviews | General | Schedule | Requirement | V1M2 4.15.3 | Management review shall be completed on an annual basis.
368 | 4.16 Data Integrity | Data Integrity |  |  | V1M2 4.16 | Data Integrity Investigations
369 | 4.16 Data Integrity | Investigation | Confidentiality | Requirement | V1M2 4.16 | All investigations resulting from data integrity issues should be conducted in a confidential manner until they are completed.
370 | 4.16 Data Integrity | Investigation | Investigation | Documentation | V1M2 4.16 | These investigations shall be documented as well as any notifications made to clients receiving any affected data.

## V1M2 Technical
ID | Major Topic | Topic | Subject | Category | Citation | Does the laboratory comply with this section? | Yes | No | N/A
370 |  |  |  |  | V1M2 5 | TECHNICAL REQUIREMENTS
371 | 5.1 General |  |  |  | V1M2 5.1.1 | General (ISO/IEC 17025:2005(E), Clause 5.1)
372 | 5.1 General | General |  |  | V1M2 5.1.1 | Many factors determine the correctness and reliability of the tests and/or calibrations performed by a laboratory. These factors include contributions from:
373 | 5.1 General | General |  |  | V1M2 5.1.1 | - human factors (5.2);
374 | 5.1 General | General |  |  | V1M2 5.1.1 | - accommodation and environmental conditions (5.3);
375 | 5.1 General | General |  |  | V1M2 5.1.1 | - test and calibration methods and method validation (5.4);
376 | 5.1 General | General |  |  | V1M2 5.1.1 | - equipment (5.5);
377 | 5.1 General | General |  |  | V1M2 5.1.1 | - measurement traceability (5.6);
378 | 5.1 General | General |  |  | V1M2 5.1.1 | - sampling (5.7);
379 | 5.1 General | General |  |  | V1M2 5.1.1 | - the handling of test and calibration items (5.8).
380 | 5.1 General | General | Factors | Requirement | V1M2 5.1.2 | The extent to which the factors contribute to the total uncertainty of measurement differs considerably between (types of) tests and between (types of) calibrations.
381 | 5.1 General | General | Factors | Requirement | V1M2 5.1.2 | The laboratory shall take account of these factors (V1M2 5.1.1) in developing test and calibration methods and procedures, in the training and qualification of personnel, and in the selection and calibration of the equipment it uses.
382 | 5.2 Personnel | Personnel |  |  | V1M2 5.2 | Personnel (ISO/IEC 17025:2005(E), Clause 5.2)
383 | 5.2 Personnel | 5.2.1 Management Responsibilities | Competency | Requirement | V1M2 5.2.1 | The laboratory management shall ensure the competence of all who operate specific equipment, perform tests and/or calibrations, evaluate results, and sign test reports and calibration certificates. (See Section 1.6 of the technical modules)
384 | 5.2 Personnel | 5.2.1 Management Responsibilities | Training | Requirement | V1M2 5.2.1 | When using staff who are undergoing training, appropriate supervision shall be provided.
385 | 5.2 Personnel | 5.2.1 Management Responsibilities | Qualifications | Requirement | V1M2 5.2.1 | Personnel performing specific tasks shall be qualified on the basis of appropriate education, training, experience and/or demonstrated skills, as required.
386 | 5.2 Personnel | 5.2.1 Management Responsibilities |  |  | V1M2 5.2.1 N1 | NOTE 1: In some technical areas (e.g. non-destructive testing) it may be required that the personnel performing certain tasks hold personnel certification. The laboratory is responsible for fulfilling specified personnel certification requirements. The requirements for personnel certification might be regulatory, included in the standards for the specific technical field, or required by the customer.
387 | 5.2 Personnel | 5.2.1 Management Responsibilities |  |  | V1M2 5.2.1 N2 | NOTE 2: The personnel responsible for the opinions and interpretation included in test reports should, in addition to the appropriate qualifications, training, experience and satisfactory knowledge of the testing carried out, also have:
388 | 5.2 Personnel | 5.2.1 Management Responsibilities |  |  | V1M2 5.2.1 N2 | — relevant knowledge of the technology used for the manufacturing of the items, materials, products, etc. tested, or the way they are used or intended to be used, and of the defects or degradations which may occur during or in service;
389 | 5.2 Personnel | 5.2.1 Management Responsibilities |  |  | V1M2 5.2.1 N2 | — knowledge of the general requirements expressed in the legislation and standards; and
390 | 5.2 Personnel | 5.2.1 Management Responsibilities |  |  | V1M2 5.2.1 N2 | — an understanding of the significance of deviations found with regard to the normal use of the items, materials, products, etc. concerned.
391 | 5.2 Personnel | 5.2.2 Training | Goals | Requirement | V1M2 5.2.2 | The management of the laboratory shall formulate the goals with respect to the education, training and skills of the laboratory personnel.
392 | 5.2 Personnel | 5.2.2 Training | Training Needs | Procedure | V1M2 5.2.2 | The laboratory shall have a policy and procedures for identifying training needs and providing training of personnel.
393 | 5.2 Personnel | 5.2.2 Training | Training Program | Requirement | V1M2 5.2.2 | The training programme shall be relevant to the present and anticipated tasks of the laboratory.
394 | 5.2 Personnel | 5.2.2 Training | Effectiveness | Requirement | V1M2 5.2.2 | The effectiveness of the training actions taken shall be evaluated.
395 | 5.2 Personnel | 5.2.3 Personnel Status | Scope | Requirement | V1M2 5.2.3 | The laboratory shall use personnel who are employed by, or under contract to, the laboratory.
396 | 5.2 Personnel | 5.2.3 Personnel Status | Expectations | Requirement | V1M2 5.2.3 | Where contracted and additional technical and key support personnel are used, the laboratory shall ensure that such personnel are supervised and competent and that they work in accordance with the laboratory's management system.
397 | 5.2 Personnel | 5.2.4 Job Descriptions | Scope | Record | V1M2 5.2.4 | The laboratory shall maintain current job descriptions for managerial, technical and key support personnel involved in tests and/or calibrations.
398 | 5.2 Personnel | 5.2.4 Job Descriptions | Job Descriptions |  | V1M2 5.2.4 N | NOTE: Job descriptions can be defined in many ways. As a minimum, the following should be defined:
399 | 5.2 Personnel | 5.2.4 Job Descriptions | Job Descriptions |  | V1M2 5.2.4 N | — the responsibilities with respect to performing tests and/or calibrations;
400 | 5.2 Personnel | 5.2.4 Job Descriptions | Job Descriptions |  | V1M2 5.2.4 N | — the responsibilities with respect to the planning of tests and/or calibrations and evaluation of results;
401 | 5.2 Personnel | 5.2.4 Job Descriptions | Job Descriptions |  | V1M2 5.2.4 N | — the responsibilities for reporting opinions and interpretations;
402 | 5.2 Personnel | 5.2.4 Job Descriptions | Job Descriptions |  | V1M2 5.2.4 N | — the responsibilities with respect to method modification and development and validation of new methods;
403 | 5.2 Personnel | 5.2.4 Job Descriptions | Job Descriptions |  | V1M2 5.2.4 N | — expertise and experience required;
404 | 5.2 Personnel | 5.2.4 Job Descriptions | Job Descriptions |  | V1M2 5.2.4 N | — qualifications and training programmes;
405 | 5.2 Personnel | 5.2.4 Job Descriptions | Job Descriptions |  | V1M2 5.2.4 N | — managerial duties.
406 | 5.2 Personnel | 5.2.5 Management Responsibilities | Delegating Responsibility | Requirement | V1M2 5.2.5 | The management shall authorize specific personnel to perform particular types of sampling, test and/or calibration, to issue test reports and calibration certificates, to give opinions and interpretations and to operate particular types of equipment.
407 | 5.2 Personnel | 5.2.5 Management Responsibilities | Personnel Records | Record | V1M2 5.2.5 | The laboratory shall maintain records of the relevant authorization(s), competence, educational and professional qualifications, training, skills and experience of all technical personnel, including contracted personnel.
408 | 5.2 Personnel | 5.2.5 Management Responsibilities | Record Accessibility | Requirement | V1M2 5.2.5 | This information shall be readily available and shall include the date on which authorization and/or competence is confirmed.
409 | 5.2 Personnel | Personnel |  |  | V1M2 5.2.5 N | NOTE: All references to Calibration Certificates in ISO/IEC 17025:2005(E) are not applicable to environmental testing.
410 | 5.2 Personnel | 5.2.6 Additional Requirements |  |  | V1M2 5.2.6 | Additional Personnel Requirements
411 | 5.2 Personnel | 5.2.6.1 Technical Manager | Technical Manager |  | V1M2 5.2.6.1 | Technical Manager Qualifications
412 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials |  | V1M2 5.2.6.1 | The applicable requirements for technical managers are given below.
413 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 a) | Any technical manager of an accredited environmental laboratory engaged in chemical analysis shall be a person with a bachelor’s degree in the chemical, environmental, biological sciences, physical sciences or engineering,
414 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 a) | Any technical manager of an accredited environmental laboratory engaged in chemical analysis shall have at least twenty-four college (24) semester credit hours in chemistry
415 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 a) | Any technical manager of an accredited environmental laboratory engaged in chemical analysis shall have at least two (2) years of experience in the environmental analysis of representative inorganic and organic analytes for which the laboratory seeks or maintains accreditation. A master’s or doctoral degree in one of the above disciplines may be substituted for one year of experience.
416 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 b) | Any technical manager of an accredited environmental laboratory limited to inorganic chemical analysis, other than metals analysis, shall be a person with at least an earned associate's degree in the chemical, physical or environmental sciences, or two (2) years of equivalent and successful college education,
417 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 b) | Any technical manager of an accredited environmental laboratory limited to inorganic chemical analysis, other than metals analysis, shall have a minimum of sixteen (16) college semester credit hours in chemistry.
418 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 b) | Any technical manager of an accredited environmental laboratory limited to inorganic chemical analysis, other than metals analysis, shall have at least two (2) years of experience performing such analysis.
419 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 c) | Any technical manager of an accredited environmental laboratory engaged in microbiological or biological analysis shall be a person with a bachelor’s degree in microbiology, biology, chemistry, environmental sciences, physical sciences or engineering
420 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 c) | Any technical manager of an accredited environmental laboratory engaged in microbiological or biological analysis shall have a minimum of sixteen college semester credit hours in general microbiology and biology
421 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 c) | Any technical manager of an accredited environmental laboratory engaged in microbiological or biological analysis shall have at least two (2) years of experience in the environmental analysis of representative analytes for which the laboratory seeks or maintains accreditation. A master’s or doctoral degree in one of the above disciplines may be substituted for one (1) year of experience.
422 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 c) | A person with an associate's degree in an appropriate field of the sciences or applied sciences, with a minimum of four (4) college semester credit hours in general microbiology may be the technical manager(s) of a laboratory engaged in microbiological analysis limited to fecal coliform, total coliform, E. coli, and standard plate count. Two (2) years of equivalent and successful college education, including the microbiology requirement, may be substituted for the associate's degree.
423 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 c) | the technical manager(s) of a laboratory engaged in microbiological analysis limited to fecal coliform, total coliform, E. coli, and standard plate count shall have one (1) year of experience in microbiological analyses.
424 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 d) | Any technical manager of an accredited environmental laboratory engaged in radiological analysis shall be a person with a bachelor’s degree in chemistry, environmental, biological sciences, physical sciences or engineering
425 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 d) | Any technical manager of an accredited environmental laboratory engaged in radiological analysis shall have twenty-four (24) college semester credit hours of chemistry
426 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 d) | Any technical manager of an accredited environmental laboratory engaged in radiological analysis shall have two (2) or more years of experience in the radiological analysis of environmental samples. A master’s or doctoral degree in one of the above disciplines may be substituted for one (1) year experience.
427 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 e) | The technical manager(s) of an accredited environmental laboratory engaged in microscopic examination of asbestos and/or airborne fibers shall meet the following requirements:
428 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 e) i | The technical manager(s) of an accredited environmental laboratory engaged in microscopic examination of asbestos and/or airborne fibers requiring the use of a transmission electron microscope shall have a bachelor's degree
429 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 e) i | The technical manager(s) of an accredited environmental laboratory engaged in microscopic examination of asbestos and/or airborne fibers requiring the use of a transmission electron microscope shall have successfully completed of courses in the use of the instrument
430 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 e) i | The technical manager(s) of an accredited environmental laboratory engaged in microscopic examination of asbestos and/or airborne fibers requiring the use of a transmission electron microscope shall have one (1) year of experience, under supervision, in the use of the instrument. Such experience shall include the identification of minerals.
431 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 e) ii | The technical manager(s) of an accredited environmental laboratory engaged in microscopic examination of asbestos and/or airborne fibers requiring the use of a polarized light microscope, shall have an associate's degree or two (2) years of college study
432 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 e) ii | The technical manager(s) of an accredited environmental laboratory engaged in microscopic examination of asbestos and/or airborne fibers requiring the use of a polarized light microscope, shall have successfully completed formal coursework in polarized light microscopy
433 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 e) ii | The technical manager(s) of an accredited environmental laboratory engaged in microscopic examination of asbestos and/or airborne fibers requiring the use of a polarized light microscope, shall have one year of experience, under supervision, in the use of the instrument. Such experience shall include the identification of minerals.
434 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 e) iii | The technical manager(s) of an accredited environmental laboratory engaged in microscopic examination of asbestos and/or airborne fibers requiring the use of a phase contrast microscope, as in the determination of airborne fibers, shall have an associate's degree or two (2) years of college study
435 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 e) iii | The technical manager(s) of an accredited environmental laboratory engaged in microscopic examination of asbestos and/or airborne fibers requiring the use of a phase contrast microscope, as in the determination of airborne fibers, shall have documentation of successful completion of formal coursework in phase contrast microscopy
436 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 e) iii | The technical manager(s) of an accredited environmental laboratory engaged in microscopic examination of asbestos and/or airborne fibers requiring the use of a phase contrast microscope, as in the determination of airborne fibers, shall have one (1) year of experience, under supervision, in the use of the instrument.
437 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 f) | Any technical manager of an accredited environmental laboratory engaged in the examination of radon in air shall have at least an associate’s degree or two (2) years of college
438 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 f) | Any technical manager of an accredited environmental laboratory engaged in the examination of radon in air shall have at least one (1) year of experience in radiation measurements
439 | 5.2 Personnel | 5.2.6.1 Technical Manager | Credentials | Requirement | V1M2 5.2.6.1 f) | Any technical manager of an accredited environmental laboratory engaged in the examination of radon in air shall have at least one (1) year of experience in the measurement of radon and/or radon progeny.
440 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Technical Manager Exceptions | Exceptions | V1M2 5.2.6.2 | Technical Manager Qualification Exceptions
441 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Credentials | Requirement | V1M2 5.2.6.2 a) | Notwithstanding any other provision of this Section, a full-time employee of a drinking water or sewage treatment facility who holds a valid treatment plant operator's certificate appropriate to the nature and size of such facility shall be deemed to meet the educational requirements as the technical manager.
442 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Credentials | Requirement | V1M2 5.2.6.2 a) | A technical manager shall have two (2) year testing experience devoted exclusively to the testing of environmental samples specified in the scope of the facility’s regulatory permit.
443 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Credentials | Requirement | V1M2 5.2.6.2 a) | Such accreditation for a water treatment facility and/or a sewage treatment facility shall be limited to the scope of that facility’s regulatory permit.
444 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Credentials | Requirement | V1M2 5.2.6.2 b) | A full-time employee of an industrial waste treatment facility with a minimum of two (2) years of experience under supervision in testing of environmental samples taken within such facility for the scope of that facility’s regulatory permit shall be deemed to meet the requirements for serving as the technical manager of an accredited laboratory.
445 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Credentials | Requirement | V1M2 5.2.6.2 b) | Such accreditation for an industrial waste treatment facility shall be limited to the scope of that facility’s regulatory permit.
446 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Credentials | Requirement | V1M2 5.2.6.2 c) | Persons who do not meet the education credential requirements but possess the requisite experience of 5.2.6.1 shall qualify as technical manager(s) subject to the following conditions.
447 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Credentials | Requirement | V1M2 5.2.6.2 c) i | Persons who do not meet the education credential requirements but possess the requisite experience of 5.2.6.1 shall qualify as technical manager(s) if the person shall be a technical manager of the laboratory on the date the laboratory applies for accreditation and/or becomes subject to accreditation under this Standard
448 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Credentials | Requirement | V1M2 5.2.6.2 c) i | Persons who do not meet the education credential requirements but possess the requisite experience of 5.2.6.1 shall qualify as technical manager(s) if the person shall have been a technical manager in that laboratory continuously for the previous twelve (12) months or more.
449 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Credentials | Requirement | V1M2 5.2.6.2 c) ii | Persons who do not meet the education credential requirements but possess the requisite experience of 5.2.6.1 shall be approved as a technical manager for only those fields of accreditation for which he/she has been technical manager in that laboratory for the previous twelve (12) months or more.
450 | 5.2 Personnel | 5.2.6.1 Technical Manager Exceptions | Grandfather Clause | Exceptions | V1M2 5.2.6.2 c) iii | A person who is admitted as a technical manager under these conditions, and leaves the laboratory, will be eligible for hire as a technical manager for the same fields of accreditation in another accredited laboratory.
451 | 5.2 Personnel | 5.2.7 Data Integrity Training |  |  | V1M2 5.2.7 | Data Integrity Training
452 | 5.2 Personnel | 5.2.7 Data Integrity Training | Frequency | Requirement | V1M2 5.2.7 | Data integrity training shall be provided as a formal part of new employee orientation and shall also be provided on an annual basis for all current employees.
453 | 5.2 Personnel | 5.2.7 Data Integrity Training | Consequences | Requirement | V1M2 5.2.7 | Employees are required to understand that any infractions of the laboratory data integrity procedures shall result in a detailed investigation that could lead to very serious consequences including immediate termination, debarment or civil/criminal prosecution.
454 | 5.2 Personnel | 5.2.7 Data Integrity Training | Documentation | Record | V1M2 5.2.7 | The initial data integrity training and the annual refresher training shall have a signature attendance sheet or other form of documentation that demonstrates all staff have participated and understand their obligations related to data integrity.
455 | 5.2 Personnel | 5.2.7 Data Integrity Training | Topic | Requirement | V1M2 5.2.7 | Data integrity training requires emphasis on the importance of proper written narration on the part of the analyst with respect to those cases where analytical data may be useful, but are in one sense or another partially deficient
456 | 5.2 Personnel | 5.2.7 Data Integrity Training | Documentation | Record | V1M2 5.2.7 | The topics covered in such training shall be documented in writing (such as an agenda) and provided to all trainees.
457 | 5.2 Personnel | 5.2.7 Data Integrity Training | Training |  | V1M2 5.2.7 | At a minimum, the following topics and activities shall be included:
458 | 5.2 Personnel | 5.2.7 Data Integrity Training | Topic | Requirement | V1M2 5.2.7.a | organizational mission and its relationship to the critical need for honesty and full disclosure in all analytical reporting, how and when to report data integrity issues, and record keeping;
459 | 5.2 Personnel | 5.2.7 Data Integrity Training | Topic | Requirement | V1M2 5.2.7.b | training, including discussion regarding all data integrity procedures;
460 | 5.2 Personnel | 5.2.7 Data Integrity Training | Topic | Requirement | V1M2 5.2.7.c | data integrity training documentation;
461 | 5.2 Personnel | 5.2.7 Data Integrity Training | Topic | Requirement | V1M2 5.2.7.d | in-depth data monitoring and data integrity procedure documentation; and
462 | 5.2 Personnel | 5.2.7 Data Integrity Training | Topic | Requirement | V1M2 5.2.7.e | specific examples of breaches of ethical behavior such as improper data manipulations, adjustments of instrument time clocks, and inappropriate changes in concentrations of standards.
463 | 5.2 Personnel | 5.2.7 Data Integrity Training | Topic | Requirement | V1M2 5.2.7 | The data integrity procedures may also include written ethics agreements, examples of improper practices, examples of improper chromatographic manipulations, requirements for external ethics program training, and any external resources available to employees.
464 | 5.3 Accommodation & Environmental Conditions |  |  |  | V1M2 5.3 | Accommodation and Environmental Conditions (ISO/IEC 17025:2005(E), Clause 5.3)
465 | 5.3 Accommodation & Environmental Conditions | 5.3.1 Environmental Conditions | Scope | Requirement | V1M2 5.3.1 | Laboratory facilities for testing and/or calibration, including but not limited to energy sources, lighting and environmental conditions, shall be such as to facilitate correct performance of the tests and/or calibrations.
466 | 5.3 Accommodation & Environmental Conditions | 5.3.1 Environmental Conditions | Environmental Conditions | Requirement | V1M2 5.3.1 | The laboratory shall ensure that the environmental conditions do not invalidate the results or adversely affect the required quality of any measurement.
467 | 5.3 Accommodation & Environmental Conditions | 5.3.1 Environmental Conditions | Off-site | Requirement | V1M2 5.3.1 | Particular care shall be taken when sampling and tests and/or calibrations are undertaken at sites other than a permanent laboratory facility.
468 | 5.3 Accommodation & Environmental Conditions | 5.3.1 Environmental Conditions | Technical Requirements | Record | V1M2 5.3.1 | The technical requirements for accommodation and environmental conditions that can affect the results of tests and calibrations shall be documented.
469 | 5.3 Accommodation & Environmental Conditions | 5.3.2 Monitoring | Scope | Record | V1M2 5.3.2 | The laboratory shall monitor, control and record environmental conditions as required by the relevant specifications, methods and procedures or where they influence the quality of the results.
470 | 5.3 Accommodation & Environmental Conditions | 5.3.2 Monitoring | Specific Conditions | Requirement | V1M2 5.3.2 | Due attention shall be paid, for example, to biological sterility, dust, electromagnetic disturbances, radiation, humidity, electrical supply, temperature, and sound and vibration levels, as appropriate to the technical activities concerned.
471 | 5.3 Accommodation & Environmental Conditions | 5.3.2 Monitoring | Unacceptable Conditions | Requirement | V1M2 5.3.2 | Tests and calibrations shall be stopped when the environmental conditions jeopardize the results of the tests and/or calibrations.
472 | 5.3 Accommodation & Environmental Conditions | 5.3.3 Floor Plan | Incompatible Activities | Requirement | V1M2 5.3.3 | There shall be effective separation between neighbouring areas in which there are incompatible activities.
473 | 5.3 Accommodation & Environmental Conditions | 5.3.3 Floor Plan | Cross Contamination | Requirement | V1M2 5.3.3 | Measures shall be taken to prevent cross-contamination.
474 | 5.3 Accommodation & Environmental Conditions | 5.3.4 Accessibility | Controlled Access | Requirement | V1M2 5.3.4 | Access to and use of areas affecting the quality of the tests and/or calibrations shall be controlled.
475 | 5.3 Accommodation & Environmental Conditions | 5.3.4 Accessibility | Controlled Access | Requirement | V1M2 5.3.4 | The laboratory shall determine the extent of control based on its particular circumstances.
476 | 5.3 Accommodation & Environmental Conditions | 5.3.5 Housekeeping | Good Housekeeping | Requirement | V1M2 5.3.5 | Measures shall be taken to ensure good housekeeping in the laboratory.
477 | 5.3 Accommodation & Environmental Conditions | 5.3.5 Housekeeping | Good Housekeeping | Procedure | V1M2 5.3.5 | Special procedures shall be prepared where necessary.
478 | 5.4 Environmental Test Methods & Method Validation |  |  |  | V1M2 5.4 | Environmental Test Methods and Method Validation
479 | 5.4 Environmental Test Methods & Method Validation |  |  | Note | V1M2 5.4 N | NOTE: All references to Calibration Laboratories and Calibration Methods in ISO/IEC 17025:2005(E) in these Clauses are not applicable to environmental testing.
480 | 5.4 Environmental Test Methods & Method Validation | 5.4.1 General | General |  | V1M2 5.4.1 | General (ISO/IEC 17025:2005(E), Clause 5.4.1)
481 | 5.4 Environmental Test Methods & Method Validation | 5.4.1 General | Scope | Procedure | V1M2 5.4.1 | The laboratory shall use appropriate methods and procedures for all tests and/or calibrations within its scope.
482 | 5.4 Environmental Test Methods & Method Validation | 5.4.1 General | Scope | Procedure | V1M2 5.4.1 | These include sampling, handling, transport, storage and preparation of items to be tested and/or calibrated, and, where appropriate, an estimation of the measurement uncertainty as well as statistical techniques for analysis of test and/or calibration data.
483 | 5.4 Environmental Test Methods & Method Validation | 5.4.1 General | Equipment Use | Procedure | V1M2 5.4.1 | The laboratory shall have instructions on the use and operation of all relevant equipment, and on the handling and preparation of items for testing and/or calibration, or both, where the absence of such instructions could jeopardize the results of tests and/or calibrations.
484 | 5.4 Environmental Test Methods & Method Validation | 5.4.1 General | Manual | Requirement | V1M2 5.4.1 | All instructions, standards, manuals and reference data relevant to the work of the laboratory shall be kept up to date and shall be made readily available to personnel (see 4.3).
485 | 5.4 Environmental Test Methods & Method Validation | 5.4.1 General | Deviations | Record | V1M2 5.4.1 | Deviation from test and calibration methods shall occur only if the deviation has been documented, technically justified, authorized, and accepted by the customer.
486 | 5.4 Environmental Test Methods & Method Validation | 5.4.1 General | General |  | V1M2 5.4.1 N | NOTE: International, regional or national standards or other recognized specifications that contain sufficient and concise information on how to perform the tests and/or calibrations do not need to be supplemented or rewritten as internal procedures if these standards are written in a way that they can be used as published by the operating staff in a laboratory. It may be necessary to provide additional documentation for optional steps in the method or additional details.
487 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Method Selection |  | V1M2 5.4.2 | Selection of Methods (ISO/IEC 17025:2005(E), Clause 5.4.2)
488 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Scope | Requirement | V1M2 5.4.2 | The laboratory shall use test and/or calibration methods, including methods for sampling, which meet the needs of the customer and which are appropriate for the tests and/or calibrations it undertakes.
489 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Method Sources | Requirement | V1M2 5.4.2 | Methods published in international, regional or national standards shall preferably be used.
490 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Method Edition | Requirement | V1M2 5.4.2 | The laboratory shall ensure that it uses the latest valid edition of a standard unless it is not appropriate or possible to do so.
491 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Supplemental Information | Documentation | V1M2 5.4.2 | When necessary, the standard shall be supplemented with additional details to ensure consistent application.
492 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Appropriate Methods | Requirement | V1M2 5.4.2 | When the customer does not specify the method to be used, the laboratory shall select appropriate methods that have been published either in international, regional or national standards, or by reputable technical organizations, or in relevant scientific texts or journals, or as specified by the manufacturer of the equipment.
493 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Laboratory Method | Requirement | V1M2 5.4.2 | Laboratory-developed methods or methods adopted by the laboratory may also be used if they are appropriate for the intended use and if they are validated.
494 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Customer Communication | Requirement | V1M2 5.4.2 | The customer shall be informed as to the method chosen.
495 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Method Verification | Requirement | V1M2 5.4.2 | The laboratory shall confirm that it can properly operate standard methods before introducing the tests or calibrations.
496 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Method reverification | Requirement | V1M2 5.4.2 | If the standard method changes, the confirmation shall be repeated.
497 | 5.4 Environmental Test Methods & Method Validation | 5.4.2 Method Selection | Out-of date Methods | Requirement | V1M2 5.4.2 | The laboratory shall inform the customer when the method proposed by the customer is considered to be inappropriate or out of date.
498 | 5.4 Environmental Test Methods & Method Validation | 5.4.3 Laboratory-Developed Methods | Lab Developed Methods |  | V1M2 5.4.3 | Laboratory-Developed Methods (ISO/IEC 17025:2005(E), Clause 5.4.3)
499 | 5.4 Environmental Test Methods & Method Validation | 5.4.3 Laboratory-Developed Methods | Planned Activity | Requirement | V1M2 5.4.3 | The introduction of test and calibration methods developed by the laboratory for its own use shall be a planned activity and shall be assigned to qualified personnel equipped with adequate resources.
500 | 5.4 Environmental Test Methods & Method Validation | 5.4.3 Laboratory-Developed Methods | Communication | Requirement | V1M2 5.4.3 | Plans shall be updated as development proceeds and effective communication amongst all personnel involved shall be ensured.
501 | 5.4 Environmental Test Methods & Method Validation | 5.4.4 Non-Standard Methods | Lab Developed Methods |  | V1M2 5.4.4 | Non-Standard Methods (ISO/IEC 17025:2005(E), Clause 5.4.4) is not applicable in this module and is addressed in specific technical modules based on technology.
502 | 5.4 Environmental Test Methods & Method Validation | 5.4.4 Test Methods C144 Validation | Lab Developed Methods |  | V1M2 5.4.5 | Validation of Methods (ISO/IEC 17025:2005(E), Clause 5.4.5) is not applicable in this module and is addressed in specific technical modules based on technology.
503 | 5.4 Environmental Test Methods & Method Validation | 5.4.6 Estimation of Uncertainty | Uncertainty |  | V1M2 5.4.6 | Estimation of Analytical Uncertainty
504 | 5.4 Environmental Test Methods & Method Validation | 5.4.6 Estimation of Uncertainty | Analytical Uncertainty | Procedure | V1M2 5.4.6 | Clause 5.4.6 of the ISO/IEC/IEC 17025:2005(E) concerning calibration testing does not apply. The following requirement replaces the ISO/IEC Clause.
505 | 5.4 Environmental Test Methods & Method Validation | 5.4.6 Estimation of Uncertainty | Analytical Uncertainty | Procedure | V1M2 5.4.6 | Environmental testing laboratories shall have a procedure(s) for estimating analytical uncertainty.
506 | 5.4 Environmental Test Methods & Method Validation | 5.4.6 Estimation of Uncertainty | Analytical Uncertainty | Option | V1M2 5.4.6 | Quality control measurement data may be used to determine analytical uncertainty.
507 | 5.4 Environmental Test Methods & Method Validation | 5.4.7 Control of Data | Control of Data |  | V1M2 5.4.7 | Control of Data (ISO/IEC 17025:2005(E), Clause 5.4.7)
508 | 5.4 Environmental Test Methods & Method Validation | 5.4.7 Control of Data | Systematic Review | Requirement | V1M2 5.4.7.1 | Calculations and data transfers shall be subject to appropriate checks in a systematic manner.
509 | 5.4 Environmental Test Methods & Method Validation | 5.4.7 Control of Data Computers & Automated Equipment | Computer Requirements |  | V1M2 5.4.7.2 | When computers or automated equipment are used for the acquisition, processing, recording, reporting, storage or retrieval of test or calibration data, the laboratory shall ensure that:
510 | 5.4 Environmental Test Methods & Method Validation | 5.4.7 Control of Data Computers & Automated Equipment | Computer software | Documentation | V1M2 5.4.7.2.a | computer software developed by the user is documented in sufficient detail
511 | 5.4 Environmental Test Methods & Method Validation | 5.4.7 Control of Data Computers & Automated Equipment | Software Validation | Requirement | V1M2 5.4.7.2.a | computer software developed by the user is suitably validated as being adequate for use;
512 | 5.4 Environmental Test Methods & Method Validation | 5.4.7 Control of Data Computers & Automated Equipment | Data Protection | Procedure | V1M2 5.4.7.2.b | procedures are established and implemented for protecting the data;
513 | 5.4 Environmental Test Methods & Method Validation | 5.4.7 Control of Data Computers & Automated Equipment | Data Protection | Procedure | V1M2 5.4.7.2.b | such procedures shall include, but not be limited to, integrity and confidentiality of data entry or collection, data storage, data transmission and data processing;
514 | 5.4 Environmental Test Methods & Method Validation | 5.4.7 Control of Data Computers & Automated Equipment | Computer maintenance | Requirement | V1M2 5.4.7.2.c | computers and automated equipment are maintained to ensure proper functioning
515 | 5.4 Environmental Test Methods & Method Validation | 5.4.7 Control of Data Computers & Automated Equipment | Computer operating conditions | Requirement | V1M2 5.4.7.2.c | computers and automated equipment are provided with the environmental and operating conditions necessary to maintain the integrity of test and calibration data.
516 | 5.4 Environmental Test Methods & Method Validation | 5.4.7 Control of Data Computers & Automated Equipment | Control of Data |  | V1M2 5.4.7.2 N | NOTE: Commercial off-the-shelf software (e.g. word processing, database and statistical programmes) in general use within their designed application range may be considered to be sufficiently validated. However, laboratory software configuration/modifications should be validated as in 5.4.7.2 a).
517 | 5.5 Calibration Requirements | Calibration |  |  | V1M2 5.5 | Calibration Requirements (ISO/IEC 17025:2005(E), Clause 5.5)
518 | 5.5 Calibration Requirements | 5.5.1 Scope | Equipment | Requirement | V1M2 5.5.1 | The laboratory shall be furnished with all items of sampling, measurement and test equipment required for the correct performance of the tests and/or calibrations (including sampling, preparation of test and/or calibration items, processing and analysis of test and/or calibration data).
519 | 5.5 Calibration Requirements | 5.5.1 Scope | Rental Equipment | Requirement | V1M2 5.5.1 | In those cases where the laboratory needs to use equipment outside its permanent control, it shall ensure that the requirements of this International Standard are met.
520 | 5.5 Calibration Requirements | 5.5.2 Equipment Requirements | Specifications | Requirement | V1M2 5.5.2 | Equipment and its software used for testing, calibration and sampling shall be capable of achieving the accuracy required and shall comply with specifications relevant to the tests and/or calibrations concerned.
521 | 5.5 Calibration Requirements | 5.5.2 Equipment Requirements | Calibration Programs | Requirement | V1M2 5.5.2 | Calibration programmes shall be established for key quantities or values of the instruments where these properties have a significant effect on the results.
522 | 5.5 Calibration Requirements | 5.5.2 Equipment Requirements | Initial Check | Requirement | V1M2 5.5.2 | Before being placed into service, equipment (including that used for sampling) shall be calibrated or checked to establish that it meets the laboratory's specification requirements and complies with the relevant standard specifications.
523 | 5.5 Calibration Requirements | 5.5.2 Equipment Requirements | Calibration before Use | Requirement | V1M2 5.5.2 | It (equipment) shall be checked and/or calibrated before use (see 5.6).
524 | 5.5 Calibration Requirements | 5.5.3 Equipment Operation | Personnel | Requirement | V1M2 5.5.3 | Equipment shall be operated by authorized personnel.
525 | 5.5 Calibration Requirements | 5.5.3 Equipment Operation | SOPs/Manuals | Requirement | V1M2 5.5.3 | Up-to-date instructions on the use and maintenance of equipment (including any relevant manuals provided by the manufacturer of the equipment) shall be readily available for use by the appropriate laboratory personnel.
526 | 5.5 Calibration Requirements | 5.5.4 Identification | Unique Identifiers | Requirement | V1M2 5.5.4 | Each item of equipment and its software used for testing and calibration and significant to the result shall, when practicable, be uniquely identified.
527 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5 | Records shall be maintained of each item of equipment and its software significant to the tests and/or calibrations performed.
528 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5 | The records shall include at least the following:
529 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.a | the identity of the item of equipment
530 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.a | the identity of the software;
531 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.b | the manufacturer's name
532 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.b | type identification
533 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.b | and serial number or other unique identification;
534 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.c | checks that equipment complies with the specification (see 5.5.2);
535 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.d | the current location, where appropriate;
536 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.e | the manufacturer's instructions, if available, or reference to their location;
537 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.f | dates, results and copies of reports and certificates of all calibrations
538 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.f | dates, results and copies of reports and certificates of all adjustments
539 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.f | dates, results and copies of reports and certificates of all acceptance criteria
540 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.f | the due date of next calibration;
541 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.g | the maintenance plan, where appropriate
542 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.g | and maintenance carried out to date;
543 | 5.5 Calibration Requirements | 5.5.5 Records | Equipment | Record | V1M2 5.5.5.h | any damage, malfunction, modification or repair to the equipment.
544 | 5.5 Calibration Requirements | 5.5.6 Maintenance | Planned Maintenance | Requirement | V1M2 5.5.6 | The laboratory shall have procedures for safe handling, transport, storage, use and planned maintenance of measuring equipment to ensure proper functioning and in order to prevent contamination or deterioration.
545 | 5.5 Calibration Requirements | 5.5.6 Maintenance |  |  | V1M2 5.5.6 N | NOTE: Additional procedures may be necessary when measuring equipment is used outside the permanent laboratory for tests, calibrations or sampling.
546 | 5.5 Calibration Requirements | 5.5.7 Defective Equipment | Removal from Service | Requirement | V1M2 5.5.7 | Equipment that has been subjected to overloading or mishandling, gives suspect results, or has been shown to be defective or outside specified limits, shall be taken out of service
547 | 5.5 Calibration Requirements | 5.5.7 Defective Equipment | Identification | Requirement | V1M2 5.5.7 | It (equipment) shall be isolated to prevent its use or clearly labelled or marked as being out of service until it has been repaired and shown by calibration or test to perform correctly.
548 | 5.5 Calibration Requirements | 5.5.7 Defective Equipment | Examining Consequences | Requirement | V1M2 5.5.7 | The laboratory shall examine the effect of the defect or departure from specified limits on previous tests and/or calibrations and shall institute the "Control of nonconforming work" procedure (see 4.9).
549 | 5.5 Calibration Requirements | 5.5.8 Calibration Status | Calibration Status | Requirement | V1M2 5.5.8 | Whenever practicable, all equipment under the control of the laboratory and requiring calibration shall be labelled, coded or otherwise identified to indicate the status of calibration
550 | 5.5 Calibration Requirements | 5.5.8 Calibration Status | Calibration Date | Requirement | V1M2 5.5.8 | Whenever practicable, all equipment under the control of the laboratory and requiring calibration shall be labelled, coded or otherwise identified to indicate the date when last calibrated
551 | 5.5 Calibration Requirements | 5.5.8 Calibration Status | Expiration | Requirement | V1M2 5.5.8 | Whenever practicable, all equipment under the control of the laboratory and requiring calibration shall be labelled, coded or otherwise identified to indicate the date or expiration criteria when recalibration is due.
552 | 5.5 Calibration Requirements | 5.5.9 After Equipment Returns | Function and Calibration | Record | V1M2 5.5.9 | When, for whatever reason, equipment goes outside the direct control of the laboratory, the laboratory shall ensure that the function and calibration status of the equipment are checked and shown to be satisfactory before the equipment is returned to service.
553 | 5.5 Calibration Requirements | 5.5.10 Intermediate Checks | Frequency | Procedure | V1M2 5.5.10 | When intermediate checks are needed to maintain confidence in the calibration status of the equipment, these checks shall be carried out according to a defined procedure.
554 | 5.5 Calibration Requirements | 5.5.11 Correction Factors | Updates | Procedure | V1M2 5.5.11 | Where calibrations give rise to a set of correction factors, the laboratory shall have procedures to ensure that copies (e.g. in computer software) are correctly updated.
555 | 5.5 Calibration Requirements | 5.5.12 Inadvertent Adjustments | Protection | Requirement | V1M2 5.5.12 | Test and calibration equipment, including both hardware and software, shall be safeguarded from adjustments which would invalidate the test and/or calibration results.
556 | 5.5 Calibration Requirements | Calibration |  |  | V1M2 5.5.12 N | NOTE: ISO/IEC Clauses 5.5.1 to 5.5.12 apply with respect to equipment in environmental testing laboratories.
557 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements |  |  | V1M2 5.5.13 | Additional Requirements and Clarifications
558 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements |  |  | V1M2 5.5.13 | Calibration requirements for analytical support equipment are included in this Section while requirements for instrument (testing) calibration are included in technical modules (i.e., Asbestos, Chemistry, Microbiology, Radiochemistry and Toxicology).
559 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment |  |  | V1M2 5.5.13.1 | Support Equipment
560 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Scope |  | V1M2 5.5.13.1 | This Standard applies to all devices that may not be the actual test instrument, but are necessary to support laboratory operations. These include, but are not limited to: balances, ovens, refrigerators, freezers, incubators, water baths, temperature measuring devices (including thermometers and thermistors), thermal/pressure sample preparation devices and volumetric dispensing devices (such as Eppendorf® or automatic dilutor/dispensing devices), if quantitative results are dependent on their accuracy, as in standard preparation and dispensing or dilution into a specified volume.
561 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Maintenance | Requirement | V1M2 5.5.13.1.a | All support equipment shall be maintained in proper working order.
562 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Maintenance | Record | V1M2 5.5.13.1.a | The records of all repair and maintenance activities, including service calls, shall be kept.
563 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Annual Calibration/verification | Requirement | V1M2 5.5.13.1.b | All support equipment shall be calibrated or verified at least annually, using a recognized National Metrology Institute, such as NIST, traceable references when available, bracketing the range of use.
564 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Calibration/verification Results | Requirement | V1M2 5.5.13.1.b | The results of such calibration or verification shall be within the specifications required of the application for which this equipment is used or:
565 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Calibration/verification Results | Option | V1M2 5.5.13.1.b.i | i. the equipment shall be removed from service until repaired; or
566 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Calibration/verification Results | Option | V1M2 5.5.13.1.b.ii | ii. the laboratory shall maintain records of established correction factors to correct all measurements.
567 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Raw Data | Record | V1M2 5.5.13.1.c | Raw data records shall be retained to document equipment performance.
568 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Daily Checks | Documentation | V1M2 5.5.13.1.d | On each day the equipment is used, balances, ovens, refrigerators, freezers and water baths shall be checked and documented.
569 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Daily Checks | Requirement | V1M2 5.5.13.1.d | The acceptability for use or continued use shall be according to the needs of the analysis or application for which the equipment is being used.
570 | 5.5 Calibration Requirements | 5.5.13 Additional Requirements Support Equipment | Volumetric Devices | Requirement | V1M2 5.5.13.1.e | Volumetric dispensing devices (except Class A glassware and Glass microliter syringes) shall be checked for accuracy on a quarterly basis.
571 | 5.6 Measurement Traceability | Measurement Traceability |  |  | V1M2 5.6 | Measurement Traceability
572 | 5.6 Measurement Traceability | Measurement Traceability | General |  | V1M2 5.6.1 | General (ISO/IEC 17025:2005(E), Clause 5.6.1) is not applicable to environmental testing.
573 | 5.6 Measurement Traceability | Measurement Traceability | General |  | V1M2 5.6.2 | Specific Requirements (ISO/IEC 17025:2005(E), Clause 5.6.2) is not applicable to environmental testing.
574 | 5.6 Measurement Traceability | 5.6.3 Reference Standards & Materials |  |  | V1M2 5.6.3 | Reference Standards and Reference Materials (ISO/IEC 17025:2005(E), Clause 5.6.3)
575 | 5.6 Measurement Traceability | 5.6.3.1 Reference Standards | Reference Standard |  | V1M2 5.6.3.1 | Reference Standards
576 | 5.6 Measurement Traceability | 5.6.3.1 Reference Standards | Reference Standard | Procedure | V1M2 5.6.3.1 | The laboratory shall have a programme and procedure for the calibration of its reference standards.
577 | 5.6 Measurement Traceability | 5.6.3.1 Reference Standards | Calibration | Requirement | V1M2 5.6.3.1 | Reference standards shall be calibrated by a body that can provide traceability as described in 5.6.2.1.
578 | 5.6 Measurement Traceability | 5.6.3.1 Reference Standards | Use | Requirement | V1M2 5.6.3.1 | Such reference standards of measurement held by the laboratory shall be used for calibration only and for no other purpose, unless it can be shown that their performance as reference standards would not be invalidated.
579 | 5.6 Measurement Traceability | 5.6.3.1 Reference Standards | Calibration | Requirement | V1M2 5.6.3.1 | Reference standards shall be calibrated before and after any adjustment.
580 | 5.6 Measurement Traceability | 5.6.3 Materials | Reference Materials |  | V1M2 5.6.3.2 | Reference Materials
581 | 5.6 Measurement Traceability | 5.6.3 Materials | Traceability | Requirement | V1M2 5.6.3.2 | Reference materials shall, where possible, be traceable to SI units of measurement, or to certified reference materials
582 | 5.6 Measurement Traceability | 5.6.3 Materials | Checks | Requirement | V1M2 5.6.3.2 | Internal reference materials shall be checked as far as is technically and economically practicable.
583 | 5.6 Measurement Traceability | 5.6.3 Reference Standards & Materials | Intermediate checks |  | V1M2 5.6.3.3 | Intermediate Checks
584 | 5.6 Measurement Traceability | 5.6.3 Reference Standards & Materials | Intermediate checks | Requirement | V1M2 5.6.3.3 | Checks needed to maintain confidence in the calibration status of reference, primary, transfer or working standards and reference materials shall be carried out according to defined procedures and schedules.
585 | 5.6 Measurement Traceability | 5.6.3 Reference Standards & Materials | Transport and Storage |  | V1M2 5.6.3.4 | Transport and Storage
586 | 5.6 Measurement Traceability | 5.6.3 Reference Standards & Materials | Transport and Storage | Procedure | V1M2 5.6.3.4 | The laboratory shall have procedures for safe handling, transport, storage and use of reference standards and reference materials in order to prevent contamination or deterioration and in order to protect their integrity.
587 | 5.6 Measurement Traceability | 5.6.3 Reference Standards & Materials |  |  | V1M2 5.6.3.4 N | NOTE: Additional procedures may be necessary when reference standards and reference materials are used outside the permanent laboratory for tests, calibrations or sampling.
588 | 5.6 Measurement Traceability | 5.6.4 Additional Requirements | Additional Requirements |  | V1M2 5.6.4 | Additional Requirements and Clarifications
589 | 5.6 Measurement Traceability | 5.6.4.1 Reference Standards and Reference Materials | Reference Standards & Materials |  | V1M2 5.6.4.1 | Reference Standards and Reference Materials
590 | 5.6 Measurement Traceability | 5.6.4.1 Reference Standards and Reference Materials | Reference Standards & Materials | Requirement | V1M2 5.6.4.1 | The laboratory shall provide satisfactory evidence of correlation of results, for example, by participation in a suitable program of inter-laboratory comparisons, proficiency testing, or independent analysis.
591 | 5.6 Measurement Traceability | 5.6.4.1 Reference Standards | Reference Standards |  | V1M2 5.6.4.1.a | Reference Standards
592 | 5.6 Measurement Traceability | 5.6.4.1 Reference Standards | Reference Standard Traceability | Requirement | V1M2 5.6.4.1.a | Where commercially available, this traceability shall be to a national standard of measurement.
593 | 5.6 Measurement Traceability | 5.6.4.1 Reference Materials | Reference Materials |  | V1M2 5.6.4.1.b | Reference Materials
594 | 5.6 Measurement Traceability | 5.6.4.1 Reference Materials | Reference Material Traceability | Requirement | V1M2 5.6.4.1.b | Where possible, traceability shall be to national or international standards of measurement or to national or international standard reference materials.
595 | 5.6 Measurement Traceability | 5.6.4.1 Reference Materials | Reference Material Checks | Requirement | V1M2 5.6.4.1.b | Internal reference materials shall be checked as far as is technically and economically practicable.
596 | 5.6 Measurement Traceability | 5.6.4.2 Standards, Reagents and Reference Materials (SRRM) |  |  | V1M2 5.6.4.2 | Documentation and Labeling of Standards, Reagents, and Reference Materials
597 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Purchase, receipt and storage | Procedure | V1M2 5.6.4.2 | Documented procedures shall exist for the purchase, receipt and storage of consumable materials used for the technical operations of the laboratory.
598 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Manufacturer | Record | V1M2 5.6.4.2.a | The laboratory shall retain records for all standards, reagents, reference materials, and media, including the manufacturer/vendor
599 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Certificate of Analysis | Record | V1M2 5.6.4.2.a | The laboratory shall retain records for all standards, reagents, reference materials, and media, including the manufacturer’s Certificate of Analysis or purity (if available)
600 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Receipt Date | Record | V1M2 5.6.4.2.a | The laboratory shall retain records for all standards, reagents, reference materials, and media, including the date of receipt
601 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Storage | Record | V1M2 5.6.4.2.a | The laboratory shall retain records for all standards, reagents, reference materials, and media, including recommended storage conditions.
602 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Expiration Date | Record | V1M2 5.6.4.2.b | For original containers, if an expiration date is provided by the manufacturer or vendor it shall be recorded on the container. If an expiration date is not provided by the manufacturer or vendor it is not required.
603 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Preparation | Record | V1M2 5.6.4.2.c | Records shall be maintained on standard, reference material, and reagent preparation. These records shall indicate traceability to purchased stocks or neat compounds, reference to the method of preparation, date of preparation, expiration date and preparer's initials
604 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Identification | Requirement | V1M2 5.6.4.2.d | All containers of prepared standards, reference materials, and reagents shall bear a unique identifier
605 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Expiration Date Prepared materials | Requirement | V1M2 5.6.4.2.d | All containers of prepared standards, reference materials, and reagents shall bear an expiration date.
606 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Reagent Suitability | Procedure | V1M2 5.6.4.2.e | Procedures shall be in place to ensure prepared reagents meet the requirements of the method.
607 | 5.6 Measurement Traceability | 5.6.4.2 (SRRM) | Use after Expiration Date | Requirement | V1M2 5.6.4.2.f | Standards, reference materials, and reagents shall not be used after their expiration dates unless their reliability is verified by the laboratory.
608 | 5.7 Collection of Samples | Sample Collection |  |  | V1M2 5.7 | Collection of Samples (ISO/IEC 17025:2005(E), Clause 5.7)
609 | 5.7 Collection of Samples | 5.7.1 Sampling Plans | Strategy | Procedure | V1M2 5.7.1 | The laboratory shall have a sampling plan and procedures for sampling when it carries out sampling of substances, materials or products for subsequent testing or calibration.
610 | 5.7 Collection of Samples | 5.7.1 Sampling Plans | Availability | Requirement | V1M2 5.7.1 | The sampling plan as well as the sampling procedure shall be available at the location where sampling is undertaken.
611 | 5.7 Collection of Samples | 5.7.1 Sampling Plans | Statistical Methods | Requirement | V1M2 5.7.1 | Sampling plans shall, whenever reasonable, be based on appropriate statistical methods.
612 | 5.7 Collection of Samples | 5.7.1 Sampling Plans | Controlling Factors | Requirement | V1M2 5.7.1 | The sampling process shall address the factors to be controlled to ensure the validity of the test and calibration results.
613 | 5.7 Collection of Samples | 5.7.1 Sampling Plans |  |  | V1M2 5.7.1 N1 | NOTE 1: Sampling is a defined procedure whereby a part of a substance, material or product is taken to provide for testing or calibration of a representative sample of the whole. Sampling may also be required by the appropriate specification for which the substance, material or product is to be tested or calibrated. In certain cases (e.g. forensic analysis), the sample may not be representative but is determined by availability.
614 | 5.7 Collection of Samples | 5.7.1 Sampling Plans |  |  | V1M2 5.7.1 N2 | NOTE 2: Sampling procedures should describe the selection, sampling plan, withdrawal and preparation of a sample or samples from a substance, material or product to yield the required information.
615 | 5.7 Collection of Samples | 5.7.2 Customer Requirements | Requested Deviations | Record | V1M2 5.7.2 | Where the customer requires deviations, additions or exclusions from the documented sampling procedure, these shall be recorded in detail with the appropriate sampling data
616 | 5.7 Collection of Samples | 5.7.2 Customer Requirements | Requested Deviations | Report | V1M2 5.7.2 | Where the customer requires deviations, additions or exclusions from the documented sampling procedure, these shall be included in all documents containing test and/or calibration results
617 | 5.7 Collection of Samples | 5.7.2 Customer Requirements | Requested Deviations | Requirement | V1M2 5.7.2 | Where the customer requires deviations, additions or exclusions from the documented sampling procedure, these shall be communicated to the appropriate personnel.
618 | 5.7 Collection of Samples | 5.7.3 Sampling Records | Sampling and/or Testing Event Records | Procedure | V1M2 5.7.3 | The laboratory shall have procedures for recording relevant data and operations relating to sampling that forms part of the testing or calibration that is undertaken.
619 | 5.7 Collection of Samples | 5.7.3 Sampling Records | Details to Record | Record | V1M2 5.7.3 | These records shall include the sampling procedure used, the identification of the sampler, environmental conditions (if relevant) and diagrams or other equivalent means to identify the sampling location as necessary and, if appropriate, the statistics the sampling procedures are based upon.
620 | 5.7 Collection of Samples | 5.7.4 Additional Requirements |  |  | V1M2 5.7.4 | Additional Requirements
621 | 5.7 Collection of Samples | 5.7.4 Additional Requirements | Sampling Date | Record | V1M2 5.7.4.a | Documentation shall include the date of sampling.
622 | 5.7 Collection of Samples | 5.7.4 Additional Requirements | Sampling Time | Record | V1M2 5.7.4.a | Documentation shall include the time of sampling.
623 | 5.7 Collection of Samples | 5.7.4 Additional Requirements | Deviations | Record | V1M2 5.7.4.b | Any deviations from sampling procedures shall be documented.
624 | 5.8 Handling Samples & Test Items | Handling Samples & Test Items |  |  | V1M2 5.8 | Handling Samples and Test Items (ISO/IEC 17025:2005(E), Clause 5.8)
625 | 5.8 Handling Samples & Test Items | 5.8.1 Procedures for |  | Procedure | V1M2 5.8.1 | The laboratory shall have procedures for the transportation, receipt, handling, protection, storage, retention and/or disposal of test and/or calibration items, including all provisions necessary to protect the integrity of the test or calibration item, and to protect the interests of the laboratory and the customer.
626 | 5.8 Handling Samples & Test Items | 5.8.2 Identification | Process | Requirement | V1M2 5.8.2 | The laboratory shall have a system for identifying test and/or calibration items.
627 | 5.8 Handling Samples & Test Items | 5.8.2 Identification | Retention | Requirement | V1M2 5.8.2 | The identification shall be retained throughout the life of the item in the laboratory.
628 | 5.8 Handling Samples & Test Items | 5.8.2 Identification | Unique Identifiers | Requirement | V1M2 5.8.2 | The system shall be designed and operated so as to ensure that items cannot be confused physically or when referred to in records or other documents.
629 | 5.8 Handling Samples & Test Items | 5.8.2 Identification | Sub-groups | Requirement | V1M2 5.8.2 | The system shall, if appropriate, accommodate a sub-division of groups of items and the transfer of items within and from the laboratory.
630 | 5.8 Handling Samples & Test Items | 5.8.3 Receipt Procedures | Departures | Record | V1M2 5.8.3 | Upon receipt of the test or calibration item, abnormalities or departures from normal or specified conditions, as described in the test or calibration method, shall be recorded
631 | 5.8 Handling Samples & Test Items | 5.8.3 Receipt Procedures | Departures communication | Requirement | V1M2 5.8.3 | When there is doubt as to the suitability of an item for test or calibration, or when an item does not conform to the description provided, or the test or calibration required is not specified in sufficient detail, the laboratory shall consult the customer for further instructions before proceeding and shall record the discussion.
632 | 5.8 Handling Samples & Test Items | 5.8.4 Storage | Storage Conditions | Procedure | V1M2 5.8.4 | The laboratory shall have procedures and appropriate facilities for avoiding deterioration, loss or damage to the test or calibration item during storage, handling and preparation.
633 | 5.8 Handling Samples & Test Items | 5.8.4 Storage | Instructions | Requirement | V1M2 5.8.4 | Handling instructions provided with the item shall be followed.
634 | 5.8 Handling Samples & Test Items | 5.8.4 Storage | Specific Conditions | Record | V1M2 5.8.4 | When items have to be stored or conditioned under specified environmental conditions, these conditions shall be maintained, monitored and recorded.
635 | 5.8 Handling Samples & Test Items | 5.8.4 Storage | Security | Requirement | V1M2 5.8.4 | Where a test or calibration item or a portion of an item is to be held secure, the laboratory shall have arrangements for storage and security that protect the condition and integrity of the secured items or portions concerned.
636 | 5.8 Handling Samples & Test Items | 5.8.4 Storage |  |  | V1M2 5.8.4 N1 | NOTE 1: Where test items are to be returned into service after testing, special care is required to ensure that they are not damaged or injured during the handling, testing or storing/waiting processes.
637 | 5.8 Handling Samples & Test Items | 5.8.4 Storage |  |  | V1M2 5.8.4 N2 | NOTE 2: A sampling procedure and information on storage and transport of samples, including information on sampling factors influencing the test or calibration result, should be provided to those responsible for taking and transporting the samples.
638 | 5.8 Handling Samples & Test Items | 5.8.4 Storage |  |  | V1M2 5.8.4 N3 | NOTE 3: Reasons for keeping a test or calibration item secure can be for reasons of record, safety or value, or to enable complementary tests and/or calibrations to be performed later.
639 | 5.8 Handling Samples & Test Items | 5.8.5 Additional Requirements |  |  | V1M2 5.8.5 | Additional Requirements - Documentation
640 | 5.8 Handling Samples & Test Items | 5.8.5 Additional Requirements |  |  | V1M2 5.8.5 | The following are essential to ensure the validity of the laboratory’s data.
641 | 5.8 Handling Samples & Test Items | 5.8.5 Additional Requirements Documentation | Unique Identifiers | Procedure | V1M2 5.8.5.a | The laboratory shall have a documented system for uniquely identifying the samples to be tested, to ensure that there can be no confusion regarding the identity of such samples at any time.
642 | 5.8 Handling Samples & Test Items | 5.8.5 Additional Requirements Documentation | Identifying Subsamples | Procedure | V1M2 5.8.5.a | This system shall include identification for all samples, sub-samples, preservations, sample containers, tests, and subsequent extracts and/or digestates.
643 | 5.8 Handling Samples & Test Items | 5.8.5 Additional Requirements Documentation | Link to Field ID | Requirement | V1M2 5.8.5.b | This laboratory code shall maintain an unequivocal link with the unique field ID code assigned to each sample.
644 | 5.8 Handling Samples & Test Items | 5.8.5 Additional Requirements Documentation | ID Placement | Requirement | V1M2 5.8.5.c | The laboratory ID code shall be placed as a durable mark on the sample container.
645 | 5.8 Handling Samples & Test Items | 5.8.5 Additional Requirements Documentation | ID Records | Record | V1M2 5.8.5.d | The laboratory ID code shall be entered into the laboratory records
646 | 5.8 Handling Samples & Test Items | 5.8.5 Additional Requirements Documentation | ID Relationships | Requirement | V1M2 5.8.5.d | The laboratory ID code shall be the link that associates the sample with related laboratory activities such as sample preparation.
647 | 5.8 Handling Samples & Test Items | 5.8.5 Additional Requirements Documentation | ID/Field ID |  | V1M2 5.8.5.e | In cases where the sample collector and analyst are the same individual, or the laboratory pre-assigns numbers to sample containers, the laboratory ID code may be the same as the field ID code.
648 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy |  |  | V1M2 5.8.6 | Additional Requirements - Sample Acceptance Policy
649 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Policy | Requirement | V1M2 5.8.6 | The laboratory shall have a written sample acceptance policy that includes the following:
650 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Sample Id | Requirement | V1M2 5.8.6.a | proper, full, and complete documentation, which shall include sample identification
651 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Sample Location | Requirement | V1M2 5.8.6.a | proper, full, and complete documentation, which shall include the location of collection
652 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Collection Date | Requirement | V1M2 5.8.6.a | proper, full, and complete documentation, which shall include the date of collection
653 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Collection Time | Requirement | V1M2 5.8.6.a | proper, full, and complete documentation, which shall include the time of collection
654 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Collector's Name | Requirement | V1M2 5.8.6.a | proper, full, and complete documentation, which shall include collector's name
655 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Preservation | Requirement | V1M2 5.8.6.a | proper, full, and complete documentation, which shall include preservation type
656 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Sample Type | Requirement | V1M2 5.8.6.a | proper, full, and complete documentation, which shall include sample type
657 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Special Remarks | Requirement | V1M2 5.8.6.a | proper, full, and complete documentation, which shall include any special remarks concerning the sample;
658 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Sample Labeling | Requirement | V1M2 5.8.6.b | proper sample labeling to include unique identification
659 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Sample Labeling | Requirement | V1M2 5.8.6.b | proper sample labeling to include a labeling system for the samples with requirements concerning the durability of the labels (water resistant) and the use of indelible ink;
660 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Sample Containers | Requirement | V1M2 5.8.6.c | use of appropriate sample containers;
661 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Holding Times | Requirement | V1M2 5.8.6.d | adherence to specified holding times;
662 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Sample Volume | Requirement | V1M2 5.8.6.e | sufficient sample volume to perform the necessary tests;
663 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Sample Deviations | Procedure | V1M2 5.8.6.f | procedures to be used when samples show signs of damage, contamination or inadequate preservation; and
664 | 5.8 Handling Samples & Test Items | 5.8.6 Additional Requirements Sample Acceptance Policy | Data Qualifications | Requirement | V1M2 5.8.6.g | qualification of any data that do not meet the above requirements.
665 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures |  |  | V1M2 5.8.7 | Additional Requirements - Sample Receipt Protocols
666 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Preservation | Procedure | V1M2 5.8.7.1 | The laboratory shall implement procedures for verifying and documenting preservation.
667 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Deviations |  | V1M2 5.8.7.2 | If the sample does not meet the sample receipt acceptance criteria listed in this Standard, the laboratory shall either:
668 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Deviations | Record | V1M2 5.8.7.2.a | retain correspondence and/or records of conversations concerning the final disposition of rejected samples; or
669 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Deviations | Record | V1M2 5.8.7.2.b | fully document any decision to proceed with the analysis of samples not meeting acceptance criteria.
670 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Deviations | Requirement | V1M2 5.8.7.2.i | The condition of these samples shall be noted on the chain of custody or transmittal form and laboratory receipt documents.
671 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Deviations | Requirement | V1M2 5.8.7.2.ii | The analysis data shall be appropriately qualified on the final report.
672 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Sample Receipt Log | Record | V1M2 5.8.7.3 | The laboratory shall utilize a permanent chronological record such as a logbook or electronic database to document receipt of all sample containers.
673 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Sample Receipt Log | Record | V1M2 5.8.7.3.a | This sample receipt log shall record the following:
674 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Sample Receipt Log Contents | Requirement | V1M2 5.8.7.3.a.i | client/project name,
675 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Sample Receipt Log Contents | Requirement | V1M2 5.8.7.3.a.ii | date and time of laboratory receipt,
676 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Sample Receipt Log Contents | Requirement | V1M2 5.8.7.3.a.iii | unique laboratory ID code, and
677 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Sample Receipt Log Contents | Requirement | V1M2 5.8.7.3.a.iv | signature or initials of the person making the entries.
678 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Supplemental Information | Requirement | V1M2 5.8.7.3.b | During the login process, the following information shall be unequivocally linked to the log record or included as a part of the log.
679 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Supplemental Information Availability | Requirement | V1M2 5.8.7.3.b | If such information is recorded/documented elsewhere, the records shall be part of the laboratory's permanent records, easily retrievable upon request and readily available to individuals who will process the sample.
680 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Supplemental Information Items | Requirement | V1M2 5.8.7.3.b N | NOTE: The placement of the laboratory ID number on the sample container is not considered a permanent record.
681 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Supplemental Information Items | Requirement | V1M2 5.8.7.3.b.i | The field ID code, which identifies each sample, shall be linked to the laboratory ID code in the sample receipt log.
682 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Supplemental Information Items | Requirement | V1M2 5.8.7.3.b.ii | The date and time of sample collection shall be linked to the sample and to the date and time of receipt in the laboratory.
683 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Supplemental Information Items | Requirement | V1M2 5.8.7.3.b.iii | The requested analyses (including applicable approved method numbers) shall be linked to the laboratory ID code.
684 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Supplemental Information Items | Requirement | V1M2 5.8.7.3.b.iv | Any comments resulting from inspection for sample rejection shall be linked to the laboratory ID code.
685 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Records Retention | Record | V1M2 5.8.7.4 | All documentation, such as memos, chain of custody, or transmittal forms that are transmitted to the laboratory by the sample transmitter, shall be retained.
686 | 5.8 Handling Samples & Test Items | 5.8.7 Additional Requirements Sample Receipt Procedures | Chain of Custody | Record | V1M2 5.8.7.5 | A complete chain of custody record form, if utilized, shall be maintained.
687 | 5.8 Handling Samples & Test Items | 5.8.8 Additional Requirements Legal Chain of Custody |  |  | V1M2 5.8.8 | Additional Requirements - Legal Chain of Custody Protocols
688 | 5.8 Handling Samples & Test Items | 5.8.8 Additional Requirements Legal Chain of Custody |  |  | V1M2 5.8.8 | Legal chain of custody procedures are used for evidentiary or legal purposes.
689 | 5.8 Handling Samples & Test Items | 5.8.8 Additional Requirements Legal Chain of Custody | Client Request | Requirement | V1M2 5.8.8 | If a client specifies that a sample is to be used for evidentiary purposes, then a laboratory shall have a written SOP for how that laboratory will carry out legal chain of custody.
690 | 5.8 Handling Samples & Test Items | 5.8.9 Additional Requirements Sample Storage and Disposal |  |  | V1M2 5.8.9 | Additional Requirements - Sample Storage and Disposal
691 | 5.8 Handling Samples & Test Items | 5.8.9 Additional Requirements Sample Storage and Disposal | Storage Conditions | Requirement | V1M2 5.8.9.a | Samples shall be stored according to the conditions specified by preservation protocols.
692 | 5.8 Handling Samples & Test Items | 5.8.9 Additional Requirements Sample Storage and Disposal | Temperature | Requirement | V1M2 5.8.9.a.i | Samples that require thermal preservation shall be stored under refrigeration that is +/-2°C of the specified preservation temperature unless regulatory or method specific criteria exist.
693 | 5.8 Handling Samples & Test Items | 5.8.9 Additional Requirements Sample Storage and Disposal | Temperature | Requirement | V1M2 5.8.9.a.i | For samples with a specified storage temperature of 4°C, storage at a temperature above the freezing point of water to 6°C shall be acceptable.
694 | 5.8 Handling Samples & Test Items | 5.8.9 Additional Requirements Sample Storage and Disposal | Storage Conditions | Requirement | V1M2 5.8.9.a.ii | Samples shall be stored away from all standards, reagents, and food. Samples shall be stored in such a manner to prevent cross contamination.
695 | 5.8 Handling Samples & Test Items | 5.8.9 Additional Requirements Sample Storage and Disposal | Storage Conditions for fractions | Requirement | V1M2 5.8.9.b | Sample fractions, extracts, leachates and other sample preparation products shall be stored according to Section 5.8.9 a) above or according to specifications in the method.
696 | 5.8 Handling Samples & Test Items | 5.8.9 Additional Requirements Sample Storage and Disposal | Sample Disposal | Procedure | V1M2 5.8.9.c | The laboratory shall have SOPs for the disposal of samples, digestates, leachates and extracts or other sample preparation products.
697 | 5.9 QA | 5.9 QA |  |  | V1M2 5.9 | Quality Assurance for Environmental Testing (ISO/IEC 17025:2005(E), Clause 5.9)
698 | 5.9 QA | 5.9.1 Quality Control | Quality Control | Requirement | V1M2 5.9.1 | The laboratory shall have quality control procedures for monitoring the validity of tests and calibrations undertaken.
699 | 5.9 QA | 5.9.1 Quality Control | Criteria for Records | Record | V1M2 5.9.1 | The resulting data shall be recorded in such a way that trends are detectable and, where practicable, statistical techniques shall be applied to the reviewing of the results.
700 | 5.9 QA | 5.9.1 Quality Control | Monitoring Program Elements | Requirement | V1M2 5.9.1 | This monitoring shall be planned and reviewed and may include, but not be limited to, the following:
701 | 5.9 QA | 5.9.1 Quality Control | Monitoring Program Elements | Option | V1M2 5.9.1.a | regular use of certified reference materials and/or internal quality control using secondary reference materials;
702 | 5.9 QA | 5.9.1 Quality Control | Monitoring Program Elements | Option | V1M2 5.9.1.b | participation in interlaboratory comparison or proficiency-testing programmes;
703 | 5.9 QA | 5.9.1 Quality Control | Monitoring Program Elements | Option | V1M2 5.9.1.c | replicate tests or calibrations using the same or different methods;
704 | 5.9 QA | 5.9.1 Quality Control | Monitoring Program Elements | Option | 5.9.1.d | retesting or recalibration of retained items;
705 | 5.9 QA | 5.9.1 Quality Control | Monitoring Program Elements | Option | V1M2 5.9.1.e | correlation of results for different characteristics of an item.
706 | 5.9 QA | 5.9.1 Quality Control |  |  | V1M2 5.9.1.e N | NOTE: The selected methods should be appropriate for the type and volume of the work undertaken.
707 | 5.9 QA | 5.9.2 Quality Control | Planned Actions | Requirement | V1M2 5.9.2 | Quality control data shall be analysed and, where they are found to be outside pre-defined criteria, planned action shall be taken to correct the problem and to prevent incorrect results from being reported.
708 | 5.9 QA | 5.9.3 Essential QC Procedures |  |  | V1M2 5.9.3 | Essential Quality Control Procedures
709 | 5.9 QA | 5.9.3 Essential QC Procedures |  |  | V1M2 5.9.3 | These general quality control principles shall apply, where applicable, to all testing laboratories. The manner in which they are implemented is dependent on the types of tests performed by the laboratory (i.e., asbestos, chemical, microbiological, radiological, toxicity) and are further described in Technical Modules. The standards for any given test type shall assure that the applicable principles are addressed:
710 | 5.9 QA | 5.9.3 Essential QC Procedures | Procedure | Requirement | V1M2 5.9.3.a | All laboratories shall have detailed written protocols in place to monitor the following quality controls:
711 |  |  |  |  | V1M2 | i. through viii are specified for the technologies identified in the technical modules. These should be used if the laboratory is using a technology other than those specified in Modules 3 through 7
712 | 5.9 QA | 5.9.3 Essential QC Procedures | Positive & Negative Controls | Requirement | V1M2 5.9.3.a.i | positive and negative controls (see technical modules), chemical or microbiological as applicable to the test type, to monitor tests such as blanks, matrix spikes, reference toxicants;
713 | 5.9 QA | 5.9.3 Essential QC Procedures | Variability | Requirement | V1M2 5.9.3.a.ii | tests to define the variability and/or repeatability of the laboratory results such as replicates;
714 | 5.9 QA | 5.9.3 Essential QC Procedures | Accuracy | Requirement | V1M2 5.9.3.a.iii | measures to assure the accuracy of the method including calibration and/or continuing calibrations, use of certified reference materials, proficiency test samples, or other measures;
715 | 5.9 QA | 5.9.3 Essential QC Procedures | Method Capability | Requirement | V1M2 5.9.3.a.iv | measures to evaluate method capability, such as limit of detection and limit of quantitation or range of applicability such as linearity;
716 | 5.9 QA | 5.9.3 Essential QC Procedures | Data Reduction | Requirement | V1M2 5.9.3.a.v | selection of appropriate formulae to reduce raw data to final results such as regression analysis, comparison to internal/external standard calculations, and statistical analyses;
717 | 5.9 QA | 5.9.3 Essential QC Procedures | Reagents and Standards | Requirement | V1M2 5.9.3.a.vi | selection and use of reagents and standards of appropriate quality;
718 | 5.9 QA | 5.9.3 Essential QC Procedures | Selectivity | Requirement | V1M2 5.9.3.a.vii | measures to assure the selectivity of the test for its intended purpose; and
719 | 5.9 QA | 5.9.3 Essential QC Procedures | Constant Test Conditions | Requirement | V1M2 5.9.3.a.viii | measures to assure constant and consistent test conditions (both instrumental and environmental) where required by the method such as temperature, humidity, light or specific instrument conditions.
720 | 5.9 QA | 5.9.3 Essential QC Procedures | Assessment | Requirement | V1M2 5.9.3.b | All quality control measures shall be assessed and evaluated on an on-going basis and quality control acceptance criteria shall be used.
721 | 5.9 QA | 5.9.3 Essential QC Procedures | Acceptance Criteria | Procedure | V1M2 5.9.3.c | The laboratory shall have procedures for the development of acceptance/rejection criteria where no method or regulatory criteria exist.
722 | 5.9 QA | 5.9.3 Essential QC Procedures | SOP Procedures | Requirement | V1M2 5.9.3.c | The quality control protocols specified by the laboratory’s SOP shall be followed (see Section 4.2.8.5 in this Standard).
723 | 5.9 QA | 5.9.3 Essential QC Procedures | Procedures in Technical Modules/Methods | Requirement | V1M2 5.9.3.c | The laboratory shall ensure that the essential standards outlined in Technical Modules or mandated methods or regulations (whichever are more stringent) are incorporated into their method manuals.
724 | 5.9 QA | 5.9.3 Essential QC Procedures | Stringency | Requirement | V1M2 5.9.3.c | When it is not apparent which is more stringent, the QC in the mandated method or regulations is to be followed.
725 | 5.10 Reporting Results |  |  |  | V1M2 5.1 | Reporting the Results
726 | 5.10 Reporting Results |  |  |  | V1M2 5.10 N | NOTE: All references to Calibration Certificates in ISO/IEC 17025:2005 are not applicable to environmental testing.
727 | 5.10 Reporting Results | 5.10.1 General |  |  | V1M2 5.10.1 | General (ISO/IEC 17025:2005(E), Clause 5.10.1)
728 | 5.10 Reporting Results | 5.10.1 General | Performance Measures | Requirement | V1M2 5.10.1 | The results of each test, calibration, or series of tests or calibrations carried out by the laboratory shall be reported accurately, clearly, unambiguously and objectively, and in accordance with any specific instructions in the test or calibration methods.
729 | 5.10 Reporting Results | 5.10.1 General | Performance Measures | Requirement | V1M2 5.10.1 | The results shall be reported, usually in a test report or a calibration certificate (see Note 1), and shall include all the information requested by the customer and necessary for the interpretation of the test or calibration results and all information required by the method used. This information is normally that required by 5.10.2, and 5.10.3 or 5.10.4.
730 | 5.10 Reporting Results | 5.10.1 General | Internal Customers | Exemption | V1M2 5.10.1 | In the case of tests or calibrations performed for internal customers, or in the case of a written agreement with the customer, the results may be reported in a simplified way.
731 | 5.10 Reporting Results | 5.10.1 General | Internal Customers | Requirement | V1M2 5.10.1 | Any information listed in 5.10.2 to 5.10.4 which is not reported to the customer shall be readily available in the laboratory which carried out the tests and/or calibrations.
732 | 5.10 Reporting Results | 5.10.1 General |  |  | V1M2 5.10.1 N1 | NOTE 1: Test reports and calibration certificates are sometimes called test certificates and calibration reports respectively.
733 | 5.10 Reporting Results | 5.10.1 General |  |  | V1M2 5.10.1 N2 | NOTE 2: The test reports or calibration certificates may be issued as hard copy or by electronic data transfer provided that the requirements of this International Standard are met.
734 | 5.10 Reporting Results | 5.10.2 Test Report |  |  | V1M2 5.10.2 | Test Reports and Calibration Certificates (ISO/IEC 17025:2005(E), Clause 5.10.2)
735 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2 | Each test report or calibration certificate shall include at least the following information, unless the laboratory has valid reasons for not doing so:
736 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.a | a title (e.g. "Test Report" or "Calibration Certificate");
737 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.b | the name and address of the laboratory, and the location where the tests and/or calibrations were carried out, if different from the address of the laboratory;
738 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.c | unique identification of the test report or calibration certificate (such as the serial number), and on each page an identification in order to ensure that the page is recognized as a part of the test report or calibration certificate, and a clear identification of the end of the test report or calibration certificate;
739 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.d | the name and address of the customer;
740 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.e | identification of the method used;
741 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.f | a description of, the condition of, and unambiguous identification of the item(s) tested or calibrated;
742 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.g | the date of receipt of the test or calibration item(s) where this is critical to the validity and application of the results, and the date(s) of performance of the test or calibration;
743 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.h | reference to the sampling plan and procedures used by the laboratory or other bodies where these are relevant to the validity or application of the results;
744 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.i | the test or calibration results with, where appropriate, the units of measurement;
745 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.j | the name(s), function(s) and signature(s) or equivalent identification of person(s) authorizing the test report or calibration certificate;
746 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2.k | where relevant, a statement to the effect that the results relate only to the items tested or calibrated.
747 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2 N1 | NOTE 1: Hard copies of test reports and calibration certificates should also include the page number and total number of pages.
748 | 5.10 Reporting Results | 5.10.2 Test Report | Contents | Requirement | V1M2 5.10.2 N2 | NOTE 2: It is recommended that laboratories include a statement specifying that the test report or calibration certificate shall not be reproduced except in full, without written approval of the laboratory.
749 | 5.10 Reporting Results | 5.10.3 Test Report |  |  | V1M2 5.10.3 | Test Reports (ISO/IEC 17025:2005(E), Clause 5.10.3)
750 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.1 | In addition to the requirements listed in 5.10.2, test reports shall, where necessary for the interpretation of the test results, include the following:
751 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.1.a | deviations from, additions to, or exclusions from the test method, and information on specific test conditions, such as environmental conditions;
752 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.1.b | where relevant, a statement of compliance/non-compliance with requirements and/or specifications;
753 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.1.c | where applicable, a statement on the estimated uncertainty of measurement; information on uncertainty is needed in test reports when it is relevant to the validity or application of the test results, when a customer's instruction so requires, or when the uncertainty affects compliance to a specification limit;
754 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.1.d | where appropriate and needed, opinions and interpretations (see 5.10.5);
755 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.1.e | additional information which may be required by specific methods, customers or groups of customers.
756 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.2 | In addition to the requirements listed in 5.10.2 and 5.10.3.1, test reports containing the results of sampling shall include the following, where necessary for the interpretation of test results:
757 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.2.a | the date of sampling;
758 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.2.b | unambiguous identification of the substance, material or product sampled (including the name of the manufacturer, the model or type of designation and serial numbers as appropriate);
759 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.2.c | the location of sampling, including any diagrams, sketches or photographs;
760 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.2.d | a reference to the sampling plan and procedures used;
761 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.2.e | details of any environmental conditions during sampling that may affect the interpretation of the test results;
762 | 5.10 Reporting Results | 5.10.3 Test Report | Additional Items | Requirement | V1M2 5.10.3.2.f | any standard or other specification for the sampling method or procedure, and deviations, additions to or exclusions from the specification concerned.
763 | 5.10 Reporting Results |  |  |  | V1M2 5.10.4 | Calibration Certificates (ISO/IEC 17025:2005(E), Clause 5.10.4) does not apply to environmental testing activities.
764 | 5.10 Reporting Results | 5.10.4 Opinions and Interpretations |  |  | V1M2 5.10.5 | Opinions and interpretations
765 | 5.10 Reporting Results | 5.10.4 Opinions and Interpretations | Basis of Interpretations | Requirement | V1M2 5.10.5 | When opinions and interpretations are included, the laboratory shall document the basis upon which the opinions and interpretations have been made. Opinions and interpretations shall be clearly marked as such in a test report.
766 | 5.10 Reporting Results | 5.10.4 Opinions and Interpretations |  |  | V1M2 5.10.5 N1 | NOTE 1: Opinions and interpretations should not be confused with inspections and product certifications as intended in ISO/IEC 17020 and ISO/IEC Guide 65.
767 | 5.10 Reporting Results | 5.10.4 Opinions and Interpretations |  |  | V1M2 5.10.5 N2 | NOTE 2: Opinions and interpretations included in a test report may comprise, but not be limited to, the following:
768 | 5.10 Reporting Results | 5.10.4 Opinions and Interpretations |  |  | V1M2 5.10.5 | • an opinion on the statement of compliance/noncompliance of the results with requirements;
769 | 5.10 Reporting Results | 5.10.4 Opinions and Interpretations |  |  | V1M2 5.10.5 | • fulfilment of contractual requirements;
770 | 5.10 Reporting Results | 5.10.4 Opinions and Interpretations |  |  | V1M2 5.10.5 | • recommendations on how to use the results;
771 | 5.10 Reporting Results | 5.10.4 Opinions and Interpretations |  |  | V1M2 5.10.5 | • guidance to be used for improvements.
772 | 5.10 Reporting Results | 5.10.4 Opinions and Interpretations |  |  | V1M2 5.10.5 N3 | NOTE 3: In many cases it might be appropriate to communicate the opinions and interpretations by direct dialogue with the customer. Such dialogue should be written down.
773 | 5.10 Reporting Results | 5.10.6 Subcontractor Results |  |  | V1M2 5.10.6 | Testing and calibration results obtained from subcontractors
774 | 5.10 Reporting Results | 5.10.6 Subcontractor Results | Identification | Requirement | V1M2 5.10.6 | When the test report contains results of tests performed by subcontractors, these results shall be clearly identified.
775 | 5.10 Reporting Results | 5.10.6 Subcontractor Results | Result Submission | Requirement | V1M2 5.10.6 | The subcontractor shall report the results in writing or electronically. When a calibration has been subcontracted, the laboratory performing the work shall issue the calibration certificate to the contracting laboratory.
776 | 5.10 Reporting Results | 5.10.7 Electronic Submission |  |  | V1M2 5.10.7 | Electronic transmission of results
777 | 5.10 Reporting Results | 5.10.7 Electronic Submission | Meeting Same Standard | Requirement | V1M2 5.10.7 | In the case of transmission of test or calibration results by telephone, telex, facsimile or other electronic or electromagnetic means, the requirements of this International Standard shall be met (see also 5.4.7).
778 | 5.10 Reporting Results | 5.10.8 Report Format |  |  | V1M2 5.10.8 | Format of reports and certificates
779 | 5.10 Reporting Results | 5.10.8 Report Format | Performance Measures | Requirement | V1M2 5.10.8 | The format shall be designed to accommodate each type of test or calibration carried out and to minimize the possibility of misunderstanding or misuse.
780 | 5.10 Reporting Results | 5.10.8 Report Format |  |  | V1M2 5.10.8 N1 | NOTE 1: Attention should be given to the lay-out of the test report or calibration certificate, especially with regard to the presentation of the test or calibration data and ease of assimilation by the reader.
781 | 5.10 Reporting Results | 5.10.8 Report Format |  |  | V1M2 5.10.8 N2 | NOTE 2: The headings should be standardized as far as possible.
782 | 5.10 Reporting Results | 5.10.9 Amendments |  |  | V1M2 5.10.9 | Amendments to test reports and calibration certificates
783 | 5.10 Reporting Results | 5.10.9 Amendments | Supplemental Reports | Requirement | V1M2 5.10.9 | Material amendments to a test report or calibration certificate after issue shall be made only in the form of a further document, or data transfer, which includes the statement: “Supplement to Test Report [or Calibration Certificate], serial number... [or as otherwise identified]”, or an equivalent form of wording.
784 | 5.10 Reporting Results | 5.10.9 Amendments | Performance Measures | Requirement | V1M2 5.10.9 | Such amendments shall meet all the requirements of this International Standard.
785 | 5.10 Reporting Results | 5.10.9 Amendments | Unique Identifiers | Requirement | V1M2 5.10.9 | When it is necessary to issue a complete new test report or calibration certificate, this shall be uniquely identified and shall contain a reference to the original that it replaces.
786 | 5.10 Reporting Results | 5.10.10 Exceptions |  |  | V1M2 5.10.10 | Exceptions
787 | 5.10 Reporting Results | 5.10.10 Exceptions | Regulatory Reports | Requirement | V1M2 5.10.10 | Some regulatory reporting requirements or formats, such as monthly operating reports, may not require all items listed below; however, the laboratory shall provide all the required information to their client for use in preparing such regulatory reports.
788 | 5.10 Reporting Results | 5.10.10 Exceptions | In-house Labs | Requirement | V1M2 5.10.10 | Laboratories operated solely to provide data for compliance purposes (in-house or captive laboratories) shall have all applicable information specified in Section 5.10 readily available for review by the accreditation body.
789 | 5.10 Reporting Results | 5.10.10 Exceptions | In-house Labs | Requirement | V1M2 5.10.10 | However, formal reports detailing the information are not required if:
790 | 5.10 Reporting Results | 5.10.10 Exceptions | In-house Labs | Requirement | V1M2 5.10.10.a | the in-house laboratory is itself responsible for preparing the regulatory reports; or
791 | 5.10 Reporting Results | 5.10.10 Exceptions | In-house Labs | Requirement | 5.10.10.b | the laboratory provides information to another individual within the organization for preparation of regulatory reports. The facility management shall ensure that the appropriate report items are in the report to the regulatory authority, if such information is required; or
792 | 5.10 Reporting Results | 5.10.10 Exceptions |  |  | 5.10.10.c | see Section 5.10.1, paragraph 3.
793 | 5.10 Reporting Results | 5.10.11 Additional Requirements |  |  | 5.10.11 | Additional Requirements
794 | 5.10 Reporting Results | 5.10.11 Additional Requirements | Time | Requirement | 5.10.11.a | Time of sample preparation and/or analysis if the required holding time for either activity is less than or equal to seventy-two hours.
795 | 5.10 Reporting Results | 5.10.11 Additional Requirements | Reporting Units | Requirement | 5.10.11.b | Results that are reported on a basis other than as received (e. g., dry weight).
796 | 5.10 Reporting Results | 5.10.11 Additional Requirements | Non-accredited Work | Requirement | 5.10.11.c | Any non-accredited tests shall be clearly identified as such to the client when claims of accreditation to this Standard are made in the analytical report or in the supporting electronic or hardcopy deliverables.
797 | 5.10 Reporting Results | 5.10.11 Additional Requirements | Results outside calibration range | Requirement | 5.10.11.d | Clear identification of numerical results with values outside the calibration range.

## Module 3
ID | Major Topic | Topic | Subject | Category | Citation | Does the laboratory comply with this section? | Yes | No | N/A
798 | 1.0 Asbestos |  |  |  | V1M3 1 | ASBESTOS TESTING
799 | 1.1 Introduction |  |  |  | V1M3 1.1 | Introduction
800 |  |  |  |  | V1M3 1.1 | This Standard applies to laboratories undertaking the examination of asbestos samples. This Standard is organized by analytical technique including transmission electron microscopy (TEM) for the analysis of water, wastewater, air, and bulk samples; phase contrast microscopy (PCM) for analysis of workplace air; and polarized light microscopy (PLM) for analysis of bulk samples. These procedures for asbestos analysis involve sample preparation followed by detection of asbestos.
801 | 1.2 Scope |  |  |  | V1M3 1.2 | Scope
802 |  |  |  |  | V1M3 1.2 | The essential quality control procedures applicable to asbestos measurements are included in this document. Additional quality control requirements that are specified by method, regulation or project shall be met by laboratories.
803 | 1.3 Terms & Definitions |  |  |  | V1M3 1.3 | Terms and Definitions
804 |  |  |  |  | V1M3 1.3 | The relevant definitions from TNI, Volume 1, Module 2, Section 3.0 are the preferred references. Definitions related to this document, which are used differently or do not exist in the above references are defined below.
805 | 1.4 Method Selection |  |  |  | V1M3 1.4 | Method Selection
806 | 1.4 Method Selection | Reference Method | Regulatory |  | V1M3 1.4 | A reference method is a method issued by an organization generally recognized as competent to do so. (When ISO refers to a standard method, that term is equivalent to reference method).
807 | 1.4 Method Selection | Reference Method | Regulatory |  | V1M3 1.4 | When a laboratory is required to analyze a parameter by a specified method due to a regulatory requirement, the parameter/method combination is recognized as a reference method.
808 | 1.4 Method Selection | Reference Method | QC | Requirement | V1M3 1.4 | If there is not a regulatory requirement for the parameter/method combination, the parameter/method combination is recognized as a reference method if it can be analyzed by another similar reference method of the same matrix and technology. The inclusion of the parameter in the method shall meet all required calibration requirements of the method and the quality control requirements of the method to which the parameter is being added. If no QC exists in the method, the laboratory shall adhere to the requirements outlined in the similar method.
809 | 1.4 Method Selection | Added Analytes | Identification | Requirement | V1M3 1.4 | A method that meets the above requirement shall be identified in such a way so that there is no confusion that the method has been modified.
810 | 1.4 Method Selection | Non Reference Methods | Customer Communication | Requirement | V1M3 1.4 | When it is necessary to use methods not covered by reference methods, these shall be subject to agreement with the client and shall include a clear specification of the client's requirements and the purpose of the environmental test.
811 | 1.4 Method Selection | Non Reference Methods | Validation | Requirement | V1M3 1.4 | The method developed shall have been validated appropriately before use.
812 | 1.5 Method Validation |  |  |  | V1M3 1.5 | Method Validation
813 | 1.5 Method Validation | General | Performance Measures | Requirement | V1M3 1.5 | Validation is the confirmation, by examination and objective evidence, that the particular requirements for a specific intended use are fulfilled.
814 | 1.5 Method Validation | General | Scope | Requirement | V1M3 1.5 | The laboratory shall validate non-reference methods, laboratory-designed/developed methods, reference methods used outside their published scope, and amplifications and modifications of reference methods to confirm that the methods are fit for the intended use.
815 | 1.5 Method Validation | General | Extent | Requirement | V1M3 1.5 | The validation shall be as extensive as is necessary to meet the needs of the given application or field of application.
816 | 1.5 Method Validation | General | Documentation | Record | V1M3 1.5 | The laboratory shall record the results obtained, the procedure used for the validation, and a statement as to whether the method is fit for the intended use.
817 | 1.5 Method Validation | General | PT Program | Requirement | V1M3 1.5 | Laboratories shall participate in proficiency testing programs. The results of these analyses shall be used to evaluate the ability of the laboratory to produce acceptable data.
818 | 1.6 Demonstration of Capability |  |  |  | V1M3 1.6 | Demonstration of Capability (DOC)
819 | 1.6 Demonstration of Capability | 1.6.1 General |  |  | V1M3 1.6.1 | General
820 | 1.6 Demonstration of Capability | 1.6.1 General | IDOC | Requirement | V1M3 1.6.1 | Prior to acceptance and institution of any method for data reporting, satisfactory initial DOC is required (see Section 1.6.2).
821 | 1.6 Demonstration of Capability | 1.6.1 General | Ongoing DOC | Requirement | V1M3 1.6.1 | Thereafter, ongoing DOC (Section 1.6.3), as per the quality control requirements in Section 1.7.3 (such as laboratory control samples) is required.
822 | 1.6 Demonstration of Capability | 1.6.1 General | IDOC | Exemption | V1M3 1.6.1 | In cases where a laboratory analyzes samples using a method that has been in use by the laboratory for at least one year prior to applying for accreditation, and there have been no significant changes in instrument type, personnel or method, the on-going DOC shall be acceptable as an initial DOC. The laboratory shall have records on file to demonstrate that an initial DOC is not required.
823 | 1.6 Demonstration of Capability | 1.6.1 General | IDOC | Record | V1M3 1.6.1 | For the initial DOC, appropriate records as discussed in Section 1.6.2 shall be completed. (See 1.6.2.1)
824 | 1.6 Demonstration of Capability | 1.6.1 General | Frequency | Requirement | V1M3 1.6.1 | An initial DOC shall be completed each time there is a change in instrument type, personnel, or method. (See 1.6.2)
825 | 1.6 Demonstration of Capability | 1.6.1 General | Documentation | Record | V1M3 1.6.1 | All demonstrations shall be documented. All data applicable to the demonstration shall be retained and readily available at the laboratory.
826 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC |  |  | V1M3 1.6.2 | Initial DOC
827 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Frequency | Requirement | V1M3 1.6.2 | An initial DOC shall be conducted prior to using any method, and at any time there is a change in instrument type, personnel or method or any time that a method has not been performed by the laboratory or analyst in a twelve (12) month period.
828 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Documentation | Record | V1M3 1.6.2.1 | The laboratory shall document each initial DOC in a manner such that the following information is readily available for each affected employee:
829 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Documentation Details | Requirement | V1M3 1.6.2.1.a | analyst(s) involved in preparation and/or analysis;
830 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Documentation Details | Requirement | V1M3 1.6.2.1.b | matrix;
831 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Documentation Details | Requirement | V1M3 1.6.2.1.c | analyte(s), class of analyte(s), or measured parameter(s);
832 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Documentation Details | Requirement | V1M3 1.6.2.1.d | identification of method(s) performed;
833 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Documentation Details | Requirement | V1M3 1.6.2.1.e | identification of laboratory-specific SOP used for analysis, including revision number;
834 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Documentation Details | Requirement | V1M3 1.6.2.1.f | date(s) of analysis; and
835 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Documentation Details | Requirement | V1M3 1.6.2.1.g | summary of analyses, including information outlined in Section 1.6.2.2.c.
836 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Use of Procedure | Requirement | V1M3 1.6.2.2 | For asbestos, if the method or regulation does not specify a DOC, the following procedure is acceptable. It is the responsibility of the laboratory to document that other approaches to DOC are adequate.
837 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure Details | Option | V1M3 1.6.2.2.a | The analyte(s) shall be diluted in a volume of clean quality system matrix (a sample in which no target analytes or interferences are present at concentrations that will impact the results of a specific method) sufficient to prepare four aliquots.
838 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure Details | Option | V1M3 1.6.2.2.b | At least four (4) aliquots shall be prepared and analyzed according to the method either concurrently or over a period of days.
839 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure Details | Option | V1M3 1.6.2.2.c | Using all of the results, calculate the mean recovery in the appropriate reporting units and the standard deviations of the population sample (in the same units) for each parameter of interest. When it is not possible to determine mean and standard deviations, such as for presence/absence and logarithmic values, the laboratory shall assess performance against established and documented criteria.
840 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure Details | Option | V1M3 1.6.2.2.d | Compare the information from (c) above to the corresponding acceptance criteria for precision and accuracy in the method (if applicable) or in laboratory-generated acceptance criteria (if there are not established mandatory criteria). If all parameters meet the acceptance criteria, the analysis of actual samples may begin. If any one of the parameters does not meet the acceptance criteria, the performance is unacceptable for that parameter.
841 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure Details | Option | V1M3 1.6.2.2.e | When one or more of the tested parameters fail at least one of the acceptance criteria, the analyst shall proceed according to i) or ii) below.
842 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure Details | Option | V1M3 1.6.2.2.e.i | Locate and correct the source of the problem and repeat the test for all parameters of interest beginning with c) above.
843 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure Details | Option | V1M3 1.6.2.2.e.ii | Beginning with c) above, repeat the test for all parameters that failed to meet criteria.
844 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure Details | Option | V1M3 1.6.2.2.f | Repeated failure, however, confirms a general problem with the measurement system. If this occurs, locate and correct the source of the problem and repeat the test for all compounds of interest beginning with b).
845 | 1.6 Demonstration of Capability | 1.6.3 On-Going DOC |  |  | V1M3 1.6.3 | On-Going DOC
846 | 1.6 Demonstration of Capability | 1.6.3 On-Going DOC | Use of Procedure | Requirement | V1M3 1.6.3.1 | The laboratory shall have a documented procedure describing ongoing demonstration of capability. The analyst(s) shall demonstrate on-going capability by meeting the quality control requirements of the method, laboratory SOP, client specifications, and/or this Standard. It is the responsibility of the laboratory to document that other approaches to ongoing DOC are adequate.
847 | 1.6 Demonstration of Capability | 1.6.3 On-Going DOC | Procedure |  | V1M3 1.6.3.2 | For asbestos, this ongoing DOC may be one of the following:
848 | 1.6 Demonstration of Capability | 1.6.3 On-Going DOC | Procedure | Option | V1M3 1.6.3.2.a | acceptable performance of a blind sample (single blind to the analyst);
849 | 1.6 Demonstration of Capability | 1.6.3 On-Going DOC | Procedure | Option | V1M3 1.6.3.2.a N | NOTE: Successful analysis of a blind performance sample on a similar method using the same technology (e.g., GC/MS volatiles by purge and trap for Methods 524.2, 624 or 5030/8260) would only require documentation for one of the test.
850 | 1.6 Demonstration of Capability | 1.6.3 On-Going DOC | Procedure | Option | V1M3 1.6.3.2.b | another initial DOC;
851 | 1.6 Demonstration of Capability | 1.6.3 On-Going DOC | Procedure | Option | V1M3 1.6.3.2.c | at least four (4) consecutive laboratory control samples with acceptable levels of precision and accuracy. The laboratory shall determine the acceptable limits for precision and accuracy prior to analysis. The laboratory shall tabulate or be able to readily retrieve four (4) consecutive passing laboratory control samples (LCS) for each method for each analyst each year;
852 | 1.6 Demonstration of Capability | 1.6.3 On-Going DOC | Procedure | Option | V1M3 1.6.3.2.d | a documented process of analyst review using quality control (QC) samples. QC samples can be reviewed to identify patterns for individuals or groups of analysts and determine if corrective action or retraining is necessary; or
853 | 1.6 Demonstration of Capability | 1.6.3 On-Going DOC | Procedure | Option | V1M3 1.6.3.2.e | if a) through d) are not technically feasible, then analysis of real-world samples with results within predefined acceptance criteria (as defined by the laboratory or method) shall be performed.
854 | 1.7 Technical Requirements |  |  |  | V1M3 1.7 | Technical Requirements
855 | 1.7.1 Technical Requirements Calibration | 1.7.1 Calibration |  |  | V1M3 1.7.1 | Calibration
856 | 1.7.1 Technical Requirements Calibration | 1.7.1 Calibration |  |  | V1M3 1.7.1 | Refer to methods referenced in the following Sections for specific equipment requirements. If NIST standard reference materials (SRM) specified below are unavailable, the laboratory may substitute an equivalent reference material with a certificate of analysis.
857 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM |  |  | V1M3 1.7.1.1 | Transmission Electron Microscopy
858 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM |  |  | V1M3 1.7.1.1 | Refer to methods referenced in the following Sections for specific equipment requirements.
859 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM |  |  | V1M3 1.7.1.1.1 | Water and Wastewater
860 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Conditions for Calibration | Requirement | V1M3 1.7.1.1.1 | All calibrations listed below (unless otherwise noted) shall be performed under the same analytical conditions used for routine asbestos analysis and shall be recorded in a notebook and include date and analyst’s signature. Frequencies stated below may be reduced to “before next use” if no samples are analyzed after the last calibration period has expired. Likewise, frequencies may have to be increased following non-routine maintenance or unacceptable calibration performance.
861 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Magnification | Requirement | V1M3 1.7.1.1.1.a | Magnification Calibration. Magnification calibration shall be done at the fluorescent screen, with the calibration specimen at the eucentric position, at the magnification used for fiber counting, generally 10,000 and 20,000x. A logbook shall be maintained with the dates of the calibration recorded. Calibrations shall be performed monthly to establish the stability of magnification. Calibration data shall be displayed on control charts that show trends over time.
862 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Camera Constant | Requirement | V1M3 1.7.1.1.1.b | Camera Constant. The camera length of the TEM in the Selected Area Electron Diffraction (SAED) mode shall be calibrated before SAED patterns of unknown samples are observed. The diffraction specimen shall be at the eucentric position for this calibration. This calibration shall allow accurate (<10% variation) measurement of layer-line spacings on the medium used for routine measurement, i.e., the phosphor screen or camera film. This shall also allow accurate (<5% variation) measurement of zone axis SAED patterns on permanent media (e.g., film).
863 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Frequency | Requirement | V1M3 1.7.1.1.1.b | Calibrations shall be performed monthly to establish the stability of the camera constant. Where non-asbestiform minerals may be expected (e.g., winchite, richterite, industrial talc, vermiculite, etc.), an internal camera constant standard such as gold, shall be deposited and measured on each sample to facilitate accurate indexing of zone axis SAED patterns. In such cases, layer line analysis alone shall not be used. Calibration data shall be displayed on control charts that show trends over time.
864 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Spot Size | Requirement | V1M3 1.7.1.1.1.c | Spot Size. The diameter of the smallest beam spot at crossover shall be less than 250 nm as calibrated quarterly. Calibration data shall be displayed on control charts that show trends over time.
865 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Beam Dose | Requirement | V1M3 1.7.1.1.1.d | Beam Dose. The beam dose shall be calibrated so that beam damage to chrysotile is minimized, specifically so that an electron diffraction pattern from a single fibril >1 ìm in length from a NIST SRM chrysotile sample is stable in the electron beam dose for at least 15 seconds.
866 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | EDXA |  | V1M3 1.7.1.1.1.e | Energy Dispersive X-Ray Analysis (EDXA) System
867 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | EDXA | Requirement | V1M3 1.7.1.1.1.e.i | The x-ray energy vs. channel number for the EDXA system shall be calibrated to within 20 eV for at least two peaks between 0.7 keV and 10 keV. One peak shall be from the low end (0.7 keV to 2 keV) and the other peak from the high end (7 keV to 10 keV) of this range. The calibration of the x-ray energy shall be checked prior to each analysis of samples and recalibrated if out of the specified range.
868 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Performance Measures | Requirement | V1M3 1.7.1.1.1.e.ii | the ability of the system to resolve the Na Ká line from the Cu L line shall be confirmed quarterly by obtaining a spectrum from the NIST SRM 1866 crocidolite sample on a copper grid.
869 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | K Factor | Requirement | V1M3 1.7.1.1.1.e.iii | The k-factors for elements found in asbestos (Na, Mg, Al, Si, Ca, and Fe) relative to Si shall be calibrated semiannually, or anytime the detector geometry may be altered. NIST SRM 2063a shall be used for Mg, Si, Ca, Fe, while k-factors for Na and Al may be obtained from suitable materials such as albite, kaersutite, or NIST SRM 99a. The k-factors shall be determined to a precision (2s) within 10% relative to the mean value obtained for Mg, Al, Si, Ca, and Fe, and within 20% relative to the mean value obtained for Na.
870 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | K Factor | Requirement | V1M3 1.7.1.1.1.e.iii | The k-factor relative to Si for Na shall be between 1.0 and 4.0, for Mg and Fe shall be between 1.0 and 2.0, and for Al and Ca shall be between 1.0 and 1.75. The k-factor for Mg relative to Fe shall be 1.5 or less. Calibration data shall be displayed on control charts that show trends over time.
871 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Detector Resolution | Requirement | V1M3 1.7.1.1.1.e.iv | The detector resolution shall be checked quarterly to ensure a full-width half maximum resolution of <175 eV at Mn Ká (5.90 keV). Calibration data shall be displayed on control charts that show trends over time.
872 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Abnormal X-Ray Area | Requirement | V1M3 1.7.1.1.1.e.v | The portions of a grid in a specimen holder for which abnormal x-ray spectra are generated under routine asbestos analysis conditions shall be determined and these areas shall be avoided in asbestos analysis.
873 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Sensitivity | Requirement | V1M3 1.7.1.1.1.e.vi | The sensitivity of the detector for collecting x-rays from small volumes shall be documented quarterly by collecting resolvable Mg and Si peaks from a unit fibril of NIST SRM 1866 chrysotile.
874 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Low Temperature Asher | Requirement | V1M3 1.7.1.1.1.f | Low Temperature Asher. The low temperature asher shall be calibrated quarterly by determining a calibration curve for the weight vs. ashing time of collapsed mixed-celluloseester (MCE) filters. Calibration data shall be displayed on control charts that show trends over time.
875 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Water and Wastewater | Grid Openings | Requirement | V1M3 1.7.1.1.1.g | Grid Openings. The magnification of the grid opening measurement system shall be calibrated using an appropriate standard at a frequency of 20 openings/20 grids/lot of 1000 or 1 opening/sample. The variation in the calibration measurements (2s) is <5% of the mean calibration value.
876 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Air |  |  | V1M3 1.7.1.1.2 | Air
877 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Air | Calibration | Procedure | V1M3 1.7.1.1.2 | All calibrations shall be performed in accordance with Section 1.7.1.1.1, with the exception of magnification.
878 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Air | Magnification | Requirement | V1M3 1.7.1.1.2 | Magnification calibration shall be done at the fluorescent screen, with the calibration specimen at the eucentric position, at the magnification used for fiber counting, generally 15,000 to 20,000x. A logbook shall be maintained with the dates of the calibration recorded. Calibrations shall be performed monthly to establish the stability of magnification.
879 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Air | Calibration | Record | V1M3 1.7.1.1.2 | A logbook shall be maintained with the dates of the calibration recorded.
880 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Air | Calibration | Requirement | V1M3 1.7.1.1.2 | Calibrations shall be performed monthly to establish the stability of magnification.
881 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Air |  |  | V1M3 1.7.1.1.3 | Bulk Samples
882 | 1.7.1 Technical Requirements Calibration | 1.7.1.1 Transmission EM Bulk Samples | Calibration | Procedure | V1M3 1.7.1.1.3 | All calibrations shall be performed in accordance with Section 1.7.1.1.1.
883 | 1.7.1 Technical Requirements Calibration | 1.7.1.2 Phase Contrast |  |  | V1M3 1.7.1.2 | Phase Contrast Microscopy
884 | 1.7.1 Technical Requirements Calibration | 1.7.1.2 Phase Contrast | Phase Rings | Requirement | V1M3 1.7.1.2.1 | At least once daily, the analyst shall use the telescope ocular (or Bertrand lens, for some microscopes) supplied by the manufacturer to ensure that the phase rings (annular diaphragm and phase-shifting elements) are concentric.
885 | 1.7.1 Technical Requirements Calibration | 1.7.1.2 Phase Contrast | Phase-shift detection Limit | Requirement | V1M3 1.7.1.2.2 | The phase-shift detection limit of the microscope shall be checked monthly or after modification or relocation using an HSE/NPL phase-contrast test slide for each analyst/microscope combination. This procedure assures that the minimum detectable fiber diameter (<ca. 0.25ìm) for this microscope is achieved.
886 | 1.7.1 Technical Requirements Calibration | 1.7.1.2 Phase Contrast | Counting Area | Requirement | V1M3 1.7.1.2.3 | Prior to ordering the Walton-Beckett graticule, calibration, in accordance with NIOSH 7400, Issue 2, 15 August 1994, Appendix A, shall be performed to obtain a counting area 100 ìm in diameter at the image plane. The diameter, dc (mm), of the circular counting area and the disc diameter shall be specified when ordering the graticule.
887 | 1.7.1 Technical Requirements Calibration | 1.7.1.2 Phase Contrast | Field Diameter | Requirement | V1M3 1.7.1.2.3 | The field diameter (D) shall be verified (or checked), to a tolerance of 100 ìm ± 2 ìm, with a stage micrometer upon receipt of the graticule from the manufacturer. When changes (zoom adjustment, disassembly, replacement, etc.) occur in the eyepiece-objective-reticle combination, field diameter shall be re-measured (or recalibrated) to determine field area (mm2). Recalibration of field diameter shall also be required when there is a change in interpupillary distance (i.e., change in analyst). Acceptable range for field area shall be 0.00754 mm2 to 0.00817 mm2. The actual field area shall be documented and used.
888 | 1.7.1 Technical Requirements Calibration | 1.7.1.3 Polarized Light |  |  | V1M3 1.7.1.3 | Polarized Light Microscopy
889 | 1.7.1 Technical Requirements Calibration | 1.7.1.3 Polarized Light | Microscope Alignment | Requirement | V1M3 1.7.1.3.1 | Microscope Alignment. To accurately measure the required optical properties, a properly aligned polarized light microscope (PLM) shall be utilized. The PLM shall be aligned before each use.
890 | 1.7.1 Technical Requirements Calibration | 1.7.1.3 Polarized Light | Refractive Index Liquids | Requirement | V1M3 1.7.1.3.2 | Refractive Index Liquids. Series of nD = 1.49 through 1.72 in intervals less than or equal to 0.005. Refractive index liquids for dispersion staining, high-dispersion series 1.550, 1.605, 1.680. The accurate measurement of the refractive index (RI) of a substance requires the use of calibrated refractive index liquids. These liquids shall be calibrated at first use and semiannually, or next use, whichever is less frequent, to an accuracy of 0.004, with a temperature accuracy of 2°C using a refractometer or RI glass beads.
891 | 1.7.2 Technical Requirements QC | 1.7.2 Quality Control |  |  | V1M3 1.7.2 | Quality Control
892 | 1.7.2 Technical Requirements QC | 1.7.2.1 Negative Control |  |  | V1M3 1.7.2.1 | Negative Controls
893 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM |  |  | V1M3 1.7.2.1.1 | Transmission Electron Microscopy
894 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM |  |  | V1M3 1.7.2.1.1.a | Water and Wastewater
895 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM Water & WW | Frequency | Requirement | V1M3 1.7.2.1.1.a.i | Blank determinations shall be made prior to sample collection. When using polyethylene bottles, one (1) bottle from each batch, or a minimum of one (1) from each twenty-four (24) shall be tested for background level. When using glass bottles, four (4) bottles from each twenty-four (24) shall be tested. An acceptable bottle blank level is defined as < 0.01 Million Fibers per Liter (MFL) > 10 ìm.
896 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM Water & WW | Process Blank | Requirement | V1M3 1.7.2.1.1.a.ii | A process blank sample consisting of fiber-free water shall be run before the first field sample. The quantity of water shall be > 10 mL for a 25-mm diameter filter and > 50 mL for a 47-mm diameter filter.
897 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM Air |  |  | V1M3 1.7.2.1.1.b | Air
898 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM Air | Frequency | Requirement | V1M3 1.7.2.1.1.b.i | A blank filter shall be prepared with each set of samples. A blank filter shall be left uncovered during preparation of the sample set and a wedge from that blank filter shall be prepared alongside wedges from the sample filters. At minimum, the blank filter shall be analyzed for each twenty (20) samples analyzed.
899 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM Air | Acceptance Limits | Requirement | V1M3 1.7.2.1.1.b.ii | Maximum contamination on a single blank filter shall be no more than 53 structures/mm2. Maximum average contamination for all blank filters shall be no more than 18 structures/mm2.
900 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM Bulk Samples |  |  | V1M3 1.7.2.1.1.c | Bulk Samples
901 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM Bulk Samples | Frequency | Requirement | V1M3 1.7.2.1.1.c.i | Contamination checks using asbestos-free material, such as the glass fiber blank in SRM 1866, shall be performed at a frequency of one for every twenty samples analyzed. The detection of asbestos at a concentration exceeding 0.1% will require an investigation to detect and remove the source of the asbestos contamination.
902 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM Bulk Samples | Reference Collection | Requirement | V1M3 1.7.2.1.1.c.ii | The laboratory shall maintain a list of non-asbestos fibers that can be confused with asbestos. The list shall include crystallographic and/or chemical properties that disqualify each fiber being identified as asbestos.
903 | 1.7.2 Technical Requirements QC | 1.7.2.1.1 Negative Control Transmission EM Bulk Samples | Reference Collection | Requirement | V1M3 1.7.2.1.1.c.iii | The laboratory shall have a set of reference asbestos materials, from which a set of reference diffraction and x-ray spectra may be developed.
904 | 1.7.2 Technical Requirements QC | 1.7.2.1.2 Negative Control Phase Contrast |  |  | V1M3 1.7.2.1.2 | Phase Contrast Microscopy
905 | 1.7.2 Technical Requirements QC | 1.7.2.1.2 Negative Control Phase Contrast | Field Blanks | Requirement | V1M3 1.7.2.1.2 | At least two field blanks (or 10% of the total samples, whichever is greater) shall be submitted for analysis with each set of samples. Field blanks shall be handled in a manner representative of actual handling of associated samples in the set with a single exception that air shall not be drawn through the blank sample.
906 | 1.7.2 Technical Requirements QC | 1.7.2.1.2 Negative Control Phase Contrast | Blank Cassette | Requirement | V1M3 1.7.2.1.2 | A blank cassette shall be opened for approximately thirty (30) seconds at the same time other cassettes are opened just prior to analysis.
907 | 1.7.2 Technical Requirements QC | 1.7.2.1 Negative Control Transmission EM Bulk Samples | Field Blank Results | Requirement | V1M3 1.7.2.1.2 | Results from field blank samples shall be used in the calculation to determine final airborne fiber concentration.
908 | 1.7.2 Technical Requirements QC | 1.7.2.1.2 Negative Control Phase Contrast | Blind Sample | Requirement | V1M3 1.7.2.1.2 | The identity of blank filters shall be unknown to the counter until all counts have been completed.
909 | 1.7.2 Technical Requirements QC | 1.7.2.1.2 Negative Control Phase Contrast | Field Blank Acceptance Limits | Requirement | V1M3 1.7.2.1.2 | If a field blank yields greater than seven (7) fibers per one hundred (100) graticule fields, report possible contamination of the samples.
910 | 1.7.2 Technical Requirements QC | 1.7.2.1.3 Negative Control Polarized Light |  |  | V1M3 1.7.2.1.3 | Polarized Light Microscopy
911 | 1.7.2 Technical Requirements QC | 1.7.2.1.3 Negative Control Polarized Light | Frequency | Requirement | V1M3 1.7.2.1.3.a | Friable Materials. At least one (1) blank slide shall be prepared daily or with every fifty (50) samples analyzed, whichever is less. This is prepared by mounting a sub-sample of an isotropic verified non-asbestos-containing material (non-ACM) (e.g., fiberglass in SRM 1866) in a drop of immersion oils normally used on a clean slide, rubbing preparation tools (forceps, dissecting needles, etc.) in the mount and placing a clean coverslip on the drop.
912 | 1.7.2 Technical Requirements QC | 1.7.2.1.3 Negative Control Polarized Light | Extent of Scanning | Requirement | V1M3 1.7.2.1.3.a | The entire area under the coverslip shall be scanned to detect any asbestos contamination.
913 | 1.7.2 Technical Requirements QC | 1.7.2.1.3 Negative Control Polarized Light | Frequency | Requirement | V1M3 1.7.2.1.3.a | A similar check shall be made after every twenty (20) uses of each piece of homogenization equipment.
914 | 1.7.2 Technical Requirements QC | 1.7.2.1.3 Negative Control Polarized Light | Verified non-ACM | Requirement | V1M3 1.7.2.1.3.a | An isotropic verified non-ACM shall be homogenized in the clean equipment, a slide prepared with the material and the slide scanned for asbestos contamination. (This can be substituted for the blank slide mentioned in this Section.)
915 | 1.7.2 Technical Requirements QC | 1.7.2.1.3 Negative Control Polarized Light | Frequency | Requirement | V1M3 1.7.2.1.3.b | Non-Friable Materials. At least one (1) non-ACM non-friable material shall be prepared and analyzed with every twenty (20) samples analyzed. This non-ACM shall go through the full preparation and analysis regimen for the type of analysis being performed.
916 | 1.7.3 Technical Requirements Test Variability |  |  |  | V1M3 1.7.3 | Test Variability/Reproducibility
917 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM |  |  | V1M3 1.7.3.1 | Transmission Electron Microscopy
918 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM | Frequency | Requirement | V1M3 1.7.3.1 | Quality assurance analyses shall be performed regularly covering all time periods, instruments, tasks, and personnel. The selection of samples shall be random and samples of special interest may be included in the selection of samples for quality assurance analyses. When possible, the checks on personnel performance shall be executed without their prior knowledge. A disproportionate number of analyses shall not be performed prior to internal or external audits. It is recommended that a laboratory initially be at 100% quality control (all samples re-analyzed). The proportion of quality control samples can later be lowered gradually, as control indicates, to a minimum of 10%.
919 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Water & WW |  |  | V1M3 1.7.3.1.1 | Water and Wastewater
920 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Water & WW | Relocator Grids | Requirement | V1M3 1.7.3.1.1 | All analyses shall be performed on relocator grids so that other laboratories can easily repeat analyses on the same grid openings. Quality assurance analyses shall not be postponed during periods of heavy workloads.
921 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Water & WW | Frequency | Requirement | V1M3 1.7.3.1.1 | The total number of QA samples and blanks shall be greater than or equal to 10% of the total sample workload. Precision of analyses is related to concentration, as gleaned from inter-laboratory proficiency testing. Relative standard deviations (RSD) for amphibole asbestos decreased from 50% at 0.8 MFL to 25% at 7 MFL in inter-laboratory proficiency testing, while RSD for chrysotile was higher, 50% at 6 MFL.
922 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Water & WW | Replicate Frequency | Requirement | V1M3 1.7.3.1.1.a | Replicate. A second, independent analysis shall be performed on the same grids but on different grid openings than used in the original analysis of a sample. Results shall be within 1.5x of Poisson standard deviation. This shall be performed at a frequency of one (1) per one hundred (100) samples.
923 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Water & WW | Duplicate Frequency | Requirement | V1M3 1.7.3.1.1.b | Duplicate. A second aliquot of sample shall be filtered through a second filter, prepared and analyzed in the same manner as the original preparation of that sample. Results shall be within 2.0x of Poisson standard deviation. This shall be performed at a frequency of one (1) per one hundred (100) samples.
924 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Water & WW | Verified Analysis | Requirement | V1M3 1.7.3.1.1.c | Verified Analyses. A second, independent analysis shall be performed on the same grids and grid openings used in the original analysis of a sample. The two sets of results shall be compared according to Turner and Steel (NISTIR 5351). This shall be performed at a frequency of one (1) per twenty (20) samples. Qualified analysts shall maintain an average of ≥ 80% true positives, ≤ 20% false negatives, and ≤ 10% false positives.
925 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Air |  |  | V1M3 1.7.3.1.2 | Air
926 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Air | Relocator Grids | Requirement | V1M3 1.7.3.1.2.a | All analyses shall be performed on relocator grids so that other laboratories can easily repeat analyses on the same grid openings.
927 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Air | Frequency | Requirement | V1M3 1.7.3.1.2.b | The laboratory and TEM analysts shall obtain mean analytical results on NIST SRM 1876b so that trimmed mean values fall within 80% of the lower limit and 110% of the upper limit of the 95% confidence limits as published on the certificate. These limits are derived from the allowable false positives and false negatives given in Section 1.7.3.1.1.c, Verified Analysis, below. SRM 1876b shall be analyzed a minimum of once per year by each TEM analyst.
928 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Air | Acceptance Limits | Requirement | V1M3 1.7.3.1.2.c | The laboratory shall have documentation demonstrating that TEM analysts correctly classify at least 90% of both bundles and single fibrils of asbestos structures greater than or equal to 1 μm in length in known standard materials traceable to NIST, such as NIST bulk asbestos SRM 1866.
929 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Air | Interlaboratory Analysis | Requirement | V1M3 1.7.3.1.2.d | Inter-laboratory analyses shall be performed to detect laboratory bias. The frequency of inter-laboratory verified analysis shall correspond to a minimum of one (1) per two hundred (200) grid square analyses for clients.
930 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Air | Multiple Instruments | Requirement | V1M3 1.7.3.1.2.e | If more than one TEM is used for asbestos analysis, intermicroscope analyses shall be performed to detect instrument bias.
931 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Air | Replicate Frequency | Requirement | V1M3 1.7.3.1.2.e.i | Replicate. A second, independent analysis shall be performed in accordance with Section 1.7.3.1.1.a.
932 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Air | Duplicate Frequency | Requirement | V1M3 1.7.3.1.2.e.ii | Duplicate. A second wedge from a sample filter shall be prepared and analyzed in the same manner as the original preparation of that sample. Results shall be within 2.0x of Poisson standard deviation. This shall be performed at a frequency of one (1) per one hundred (100) samples.
933 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Air | Verified Analysis | Requirement | V1M3 1.7.3.1.2.e.iii | Verified Analyses. A second, independent analysis shall be performed on the same grids and grid openings in accordance with Section 1.7.3.1.1.c.
934 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Bulk Samples |  |  | V1M3 1.7.3.1.3 | Bulk Samples
935 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Bulk Samples | Concentration (% Asbestos) | Requirement | V1M3 1.7.3.1.3 | Bulk samples with low (< 10%) asbestos content are the most problematic. At least 30% of a laboratory’s QC analyses shall be performed on samples containing from 1% to 10% asbestos.
936 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Bulk Samples | Intra-Analyst Frequency | Requirement | V1M3 1.7.3.1.3.a | Intra-Analyst Precision. At least one (1) out of fifty (50) samples shall be re-analyzed by the same analyst. For single analyst laboratories, at least one (1) out of every ten (10) samples shall be re-analyzed by the same analyst.
937 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Bulk Samples | Inter-Analyst Frequency | Requirement | V1M3 1.7.3.1.3.b | Inter-Analyst Precision. At least one (1) out of fifteen (15) samples shall be re-analyzed by another analyst. Inter-analyst results will require additional re-analysis, possibly including another analyst, to resolve discrepancies when classification (ACM vs. non-ACM) errors occur, when asbestos identification errors occur, or when inter-analyst precision is found to be unacceptable.
938 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Bulk Samples | Inter-Laboratory Frequency | Requirement | V1M3 1.7.3.1.3.c | Inter-Laboratory Precision. The laboratory shall participate in round robin testing with at least one (1) other laboratory. Samples shall be sent to this other laboratory at least four (4) times per year.
939 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Bulk Samples | Sample Type | Requirement | V1M3 1.7.3.1.3.c | These samples shall be samples previously analyzed as QC samples.
940 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Bulk Samples | Assessment | Requirement | V1M3 1.7.3.1.3.c | Results of these analyses shall be assessed in accordance with QC requirements.
941 | 1.7.3 Technical Requirements Test Variability | 1.7.3.1 Transmission EM Bulk Samples | Misclassifications or misidentification | Requirement | V1M3 1.7.3.1.3.c | The QC requirements shall address misclassifications (false positives, false negatives) and misidentification of asbestos types.
942 | 1.7.3 Technical Requirements Test Variability | 1.7.3.2 Phase Contrast |  |  | V1M3 1.7.3.2 | Phase Contrast Microscopy
943 | 1.7.3 Technical Requirements Test Variability | 1.7.3.2 Phase Contrast | Interlaboratory Analysis | Requirement | V1M3 1.7.3.2.a | Inter-Laboratory Precision. Each laboratory analyzing air samples for compliance determination shall implement an inter-laboratory quality assurance program that includes participation of at least two (2) other independent laboratories. Each laboratory shall participate in round robin testing at least once every six months with at least all the other laboratories in its inter-laboratory quality assurance group.
944 | 1.7.3 Technical Requirements Test Variability | 1.7.3.2 Phase Contrast | Sample Type | Requirement | V1M3 1.7.3.2.a | Each laboratory shall submit slides typical of its own workload for use in this program.
945 | 1.7.3 Technical Requirements Test Variability | 1.7.3.2 Phase Contrast | Statistical Analysis | Requirement | V1M3 1.7.3.2.a | The round robin shall be designed and results analyzed using appropriate statistical methodology.
946 | 1.7.3 Technical Requirements Test Variability | 1.7.3.2 Phase Contrast | Results | Requirement | V1M3 1.7.3.2.a | Results of this QA program shall be posted in each laboratory to keep the microscopists informed.
947 | 1.7.3 Technical Requirements Test Variability | 1.7.3.2 Phase Contrast | Intra and Inter Analyst | Requirement | V1M3 1.7.3.2.b | Intra- and Inter-Analyst Precision. Each analyst shall select and count a prepared slide from a “reference slide library” on each day on which air counts are performed. Reference slides shall be prepared using well-behaved samples taken from the laboratory workload. Fiber densities shall cover the entire range routinely analyzed by the laboratory. These slides shall be counted by all analysts to establish an original standard deviation and corresponding limits of acceptability. Results from the daily reference sample analysis shall be compared to the statistically derived acceptance limits using a control chart or a database. It is recommended that the labels on the reference slides be periodically changed so that the analysts do not become familiar with the samples. Intra- and inter-analyst precision may be estimated from blind recounts on reference samples. Inter-analyst precision shall be posted in each laboratory to keep the microscopists informed.
948 | 1.7.3 Technical Requirements Test Variability | 1.7.3.3 Polarized Light |  |  | V1M3 1.7.3.3 | Polarized Light Microscopy
949 | 1.7.3 Technical Requirements Test Variability | 1.7.3.3 Polarized Light | Procedure | Requirement | V1M3 1.7.3.3 | Refer to Section 1.7.3.1.3
950 | 1.7.4 Technical Requirements Other QC Measures |  |  |  | V1M3 1.7.4 | Other Quality Control Measures
951 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.1 Transmission EM |  |  | V1M3 1.7.4.1 | Transmission Electron Microscopy
952 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.1 Transmission EM Water & WW |  |  | V1M3 1.7.4.1.a | Water and Wastewater
953 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.1 Transmission EM Water & WW | Filter Preparation | Requirement | V1M3 1.7.4.1.a.i | Filter preparations shall be made from all six (6) asbestos types from NIST SRMs 1866 and 1867. These preparations shall have concentrations between one (1) and twenty (20) structures (>10μm) per 0.01 mm2. One of these preparations shall be analyzed independently at a frequency of one (1) per one hundred (100) samples analyzed. Results shall be evaluated as verified asbestos analysis in accordance with S. Turner and E.B. Steel, NISTIR 5351, Airborne Asbestos Method: Standard Test Method for Verified Analysis of Asbestos by Transmission Electron Microscopy – Version 2.0, 1994.
954 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.1 Transmission EM Water & WW | SRM | Requirement | V1M3 1.7.4.1.a.ii | NIST SRM 1876b shall be analyzed annually by each analyst. Results shall be evaluated in accordance with limits published for that SRM.
955 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.1 Transmission EM Air |  |  | V1M3 1.7.4.1.b | Air
956 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.1 Transmission EM Air | Filter Preparation | Requirement | V1M3 1.7.4.1.b.i | Filter preparations shall be made from all six (6) asbestos types in accordance with Section 1.7.4.1.a)i.
957 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.1 Transmission EM Air | SRM | Requirement | V1M3 1.7.4.1.b.ii | NIST SRM 1876b shall be analyzed annually.
958 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.1 Transmission EM Bulk Samples |  |  | V1M3 1.7.4.1.c | Bulk Samples
959 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.1 Transmission EM Bulk Samples | Identification Criterion | Requirement | V1M3 1.7.4.1.c | All analysts shall be able to correctly identify the six (6) regulated asbestos types (chrysotile, amosite, crocidolite, anthophyllite, actinolite, and tremolite). Standards for the six (6) asbestos types listed are available from NIST (SRMs 1866 and 1867).
960 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.2 Phase Contrast |  |  | V1M3 1.7.4.2 | Phase Contrast Microscopy
961 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.2 Phase Contrast | Fiber Distribution | Requirement | V1M3 1.7.4.2.a | Test for Non-Random Fiber Distribution. Blind recounts by the same analyst shall be performed on 10% of the filters counted. A person other than the counter shall re-label slides before the second count. A test for type II error shall be performed to determine whether a pair of counts by the same analyst on the same slide shall be rejected due to non-random fiber distribution. If a pair of counts is rejected by this test, the remaining samples in the set shall be recounted and the new counts shall be tested against first counts. All rejected paired counts shall be discarded.
962 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.2 Phase Contrast | Blank Recounts | Requirement | V1M3 1.7.4.2.b | It shall not be necessary to use this statistic on blank recounts.
963 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.2 Phase Contrast | National Program | Requirement | V1M3 1.7.4.2.c | All laboratories shall participate in a national sample testing scheme such as the Proficiency Analytical Testing (PAT) program or the Asbestos Analysts Registry (AAR) program, both sponsored by the American Industrial Hygiene Association (AIHA).
964 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.3 Polarized Light |  |  | V1M3 1.7.4.3 | Polarized Light Microscopy
965 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.3 Polarized Light | Friable Materials | Requirement | V1M3 1.7.4.3.a | Friable Materials. Because accuracy cannot be determined by re-analysis of routine field samples, at least one (1) out of one hundred (100) samples shall be a standard or reference sample that has been routinely resubmitted to determine analyst’s precision and accuracy. A set of these samples may be accumulated from proficiency testing samples with predetermined weight compositions or from standards generated with weighed quantities of asbestos and other bulk materials. At least half of the reference samples submitted for this QC shall contain between 1 and 10% asbestos.
966 | 1.7.4 Technical Requirements Other QC Measures | 1.7.4.3 Polarized Light | Non-friable Materials | Requirement | V1M3 1.7.4.3.b | Non-Friable Materials. At least one (1) out of one hundred (100) samples shall be a verified quantitative standard that has routinely been resubmitted to determine analyst precision and accuracy.
967 | 1.7.5 Analytical Sensitivity |  |  |  | V1M3 1.7.5 | Analytical Sensitivity
968 | 1.7.5 Analytical Sensitivity | 1.7.5.1 Transmission EM |  |  | V1M3 1.7.5.1 | Transmission Electron Microscopy
969 | 1.7.5 Analytical Sensitivity | 1.7.5.1 Transmission EM Water & WW |  |  | V1M3 1.7.5.1.1 | Water and Wastewater
970 | 1.7.5 Analytical Sensitivity | 1.7.5.1 Transmission EM Water & WW | Threshold | Requirement | V1M3 1.7.5.1.1 | An analytical sensitivity of 200,000 fibers per liter (0.2 MFL) is required for each sample analyzed. Analytical sensitivity is defined as the waterborne concentration represented by the finding of one asbestos structure in the total area of filter examined. This value will depend on the fraction of the filter sampled and the dilution factor (if applicable).
971 | 1.7.5 Analytical Sensitivity | 1.7.5.1 Transmission EM Air |  |  | V1M3 1.7.5.1.2 | Air
972 | 1.7.5 Analytical Sensitivity | 1.7.5.1 Transmission EM Air | Threshold | Requirement | V1M3 1.7.5.1.2 | An analytical sensitivity of 0.005 structures/cm2 is required for each sample analyzed. Analytical sensitivity is defined as the airborne concentration represented by the finding of one asbestos structure in the total area of filter examined. This value will depend on the effective surface area of the filter, the filter area analyzed, and the volume of air sampled.
973 | 1.7.5 Analytical Sensitivity | 1.7.5.1 Transmission EM Bulk Samples |  |  | V1M3 1.7.5.1.3 | Bulk Samples
974 | 1.7.5 Analytical Sensitivity | 1.7.5.1 Transmission EM Bulk Samples | Range | Requirement | V1M3 1.7.5.1.3 | The range is dependent on the type of bulk material being analyzed. The sensitivity may be as low as 0.0001%.
975 | 1.7.5 Analytical Sensitivity | 1.7.5.2 Phase Contrast |  |  | V1M3 1.7.5.2 | Phase Contrast Microscopy
976 | 1.7.5 Analytical Sensitivity | 1.7.5.2 Phase Contrast | Quantitative Working Range | Requirement | V1M3 1.7.5.2 | The normal quantitative working range of the method is 0.04 to 0.5 fiber/ cm2 for a 1000 L air sample. An ideal counting range on the filter shall be 100 to 1300 fibers/mm2. The limit of detection (LOD) is estimated to be 5.5 fibers per 100 fields or 7 fibers/mm2. The LOD in fiber/cc will depend on sample volume and quantity of interfering dust but shall be <0.01 fiber/cm2 for atmospheres free of interferences.
977 | 1.7.5 Analytical Sensitivity | 1.7.5.3 Polarized Light |  |  | V1M3 1.7.5.3 | Polarized Light Microscopy
978 | 1.7.5 Analytical Sensitivity | 1.7.5.3 Polarized Light | Limit of Detection | Requirement | V1M3 1.7.5.3 | The laboratory shall utilize a method that provides a limit of detection that is appropriate and relevant for the intended use of the data. Limit of detection shall be determined by the protocol in the method or applicable regulation.
979 | 1.7.6 Quality of Standards & Reagents |  |  |  | V1M3 1.7.6 | Quality of Standards and Reagents
980 | 1.7.6 Quality of Standards & Reagents | 1.7.6.1 Transmission EM |  |  | V1M3 1.7.6.1 | Transmission Electron Microscopy
981 | 1.7.6 Quality of Standards & Reagents | 1.7.6.1 Transmission EM | Need for Standards | Requirement | V1M3 1.7.6.1.a | The quality control program shall establish and maintain provisions for asbestos standards.
982 | 1.7.6 Quality of Standards & Reagents | 1.7.6.1 Transmission EM | Reference Standards | Requirement | V1M3 1.7.6.1.b | Reference standards that are used in an asbestos laboratory shall be obtained from NIST, EPA, or suppliers who participate in supplying NIST standards or NIST traceable asbestos. Any reference standards purchased outside the United States shall be traceable back to each country’s national standards laboratory. Commercial suppliers of reference standards shall conform to ANSI N42.22 to assure the quality of their products.
983 | 1.7.6 Quality of Standards & Reagents | 1.7.6.1 Transmission EM | Reference Standards | Requirement | V1M3 1.7.6.1.c | Reference standards shall be accompanied with a certificate of calibration whose content is as described in ANSI N42.22-1995, Section 8, Certificates.
984 | 1.7.6 Quality of Standards & Reagents | 1.7.6.1 Transmission EM | Reagents | Requirement | V1M3 1.7.6.1.d | All reagents used shall be analytical reagent grade or better.
985 | 1.7.6 Quality of Standards & Reagents | 1.7.6.1 Transmission EM | Mineral Fiber Collection | Requirement | V1M3 1.7.6.1.e | The laboratory shall have mineral fibers or data from mineral fibers that will allow differentiating asbestos from at least the following “look-alikes”: fibrous talc, sepiolite, wollastonite, attapulgite (palygorskite), halloysite, vermiculite scrolls, antigorite, lizardite, pyroxenes, hornblende, richterite, winchite, or any other asbestiform minerals that are suspected as being present in the sample.
986 | 1.7.6 Quality of Standards & Reagents | 1.7.6.2 Phase Contrast |  |  | V1M3 1.7.6.2 | Phase Contrast Microscopy
987 | 1.7.6 Quality of Standards & Reagents | 1.7.6.2 Phase Contrast | Standard Source | Requirement | V1M3 1.7.6.2 | Standards of known concentration have not been developed for this testing method. Routine workload samples that have been statistically validated and national proficiency testing samples such as Proficiency Analytical Testing (PAT) and Asbestos Analysts Registry (AAR) samples available from the American Industrial Hygiene Association (AIHA) may be utilized as reference samples (refer to Section D.6.2.2 b) to standardize the optical system and analyst.
988 | 1.7.6 Quality of Standards & Reagents | 1.7.6.2 Phase Contrast | Reagents | Requirement | V1M3 1.7.6.2 | All other testing reagents and devices (HSE/NPL test slide and Walton-Beckett Graticule) shall conform to the specifications of the method (refer to National Institute for Occupational Safety and Health (NIOSH) 7400, Issue 2, 15 August 1994).
989 | 1.7.6 Quality of Standards & Reagents | 1.7.6.3 Polarized Light |  |  | V1M3 1.7.6.3 | Polarized Light Microscopy
990 | 1.7.6 Quality of Standards & Reagents | 1.7.6.3 Polarized Light | Reagents and Standards | Requirement | V1M3 1.7.6.3 | Refer to Section 1.7.6.1.
991 | 1.7.7 Acceptance/Rejection Criteria |  |  |  | V1M3 1.7.7 | Data Acceptance/Rejection Criteria
992 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.1 Transmission EM |  |  | V1M3 1.7.7.1 | Transmission Electron Microscopy
993 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.1 Transmission EM Water & WW |  |  | V1M3 1.7.7.1.1 | Water and Wastewater
994 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.1 Transmission EM Water & WW | Data Reduction | Requirement | V1M3 1.7.7.1.1.a | The concentration of asbestos in a given sample shall be calculated in accordance with EPA/600/R-94/134, Method 100.2, Section 12.1.
995 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.1 Transmission EM Water & WW | Measurement Uncertainty | Requirement | V1M3 1.7.7.1.1.b | Measurement Uncertainties. The laboratory shall calculate and report the upper and lower 95% confidence limits on the mean concentration of asbestos fibers found in the sample.
996 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.1 Transmission EM Air |  |  | V1M3 1.7.7.1.2 | Air
997 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.1 Transmission EM Air | Data Reduction | Requirement | V1M3 1.7.7.1.2.a | The concentration of asbestos in a given sample shall be calculated in accordance with the method utilized.
998 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.1 Transmission EM Air | Measurement Uncertainty | Requirement | V1M3 1.7.7.1.2.b | Measurement Uncertainties. The laboratory shall calculate and report the upper and lower 95% confidence limits on the mean concentration of asbestos fibers found in the sample.
999 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.1 Transmission EM Bulk Samples |  |  | V1M3 1.7.7.1.3 | Bulk Samples
1000 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.1 Transmission EM Bulk Samples | Data Reduction | Requirement | V1M3 1.7.7.1.3.a | The concentration of asbestos in a given sample shall be calculated in accordance with the method utilized (e.g., EPA/600/R-93/116, July 1993).
1001 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.1 Transmission EM Bulk Samples | Measurement Uncertainty | Requirement | V1M3 1.7.7.1.3.b | Measurement Uncertainties. Proficiency testing for floor tiles analyzed by TEM following careful gravimetric reduction has revealed an inter-laboratory standard deviation of approximately 20% for residues containing 70% or more asbestos. Standard deviations range from 20% to 60% for residues with lower asbestos content.
1002 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.2 Phase Contrast |  |  | V1M3 1.7.7.2 | Phase Contrast Microscopy
1003 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.2 Phase Contrast | Data Reduction | Requirement | V1M3 1.7.7.2.1 | Airborne fiber concentration in a given sample shall be calculated in accordance with NIOSH 7400, Issue 2, 15 August 1994, Sections 20 and 21.
1004 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.2 Phase Contrast | Measurement Uncertainty | Requirement | V1M3 1.7.7.2.2 | Measurement Uncertainties. The laboratory shall calculate and report the intra-laboratory and inter-laboratory relative standard deviation with each set of results (NIOSH 7400, Issue 2, 15 August 1994).
1005 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.2 Phase Contrast | Reporting High Concentration | Requirement | V1M3 1.7.7.2.3 | Fiber counts above 1300 fibers/mm2 and fiber counts from samples with >50% of the filter area covered with particulate shall be reported as “uncountable” or “probably biased”. Other fiber counts outside the 100-1300 fibers/mm2 range shall be reported as having “greater than optimal variability” and as being “probably biased”.
1006 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.3 Polarized Light |  |  | V1M3 1.7.7.3 | Polarized Light Microscopy
1007 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.3 Polarized Light | Data Reduction | Requirement | V1M3 1.7.7.3.1 | The concentration of asbestos in a given sample shall be calculated in accordance with the method utilized (e.g., EPA/600/R-93/116, July 1993).
1008 | 1.7.7 Acceptance/Rejection Criteria | 1.7.7.3 Polarized Light | Measurement Uncertainty | Requirement | V1M3 1.7.7.3.2 | Method Uncertainties. Precision and accuracy shall be determined by the individual laboratory for the percent range involved. If point counting and/or visual estimates are used, a table of reasonable expanded errors shall be generated for different concentrations of asbestos.
1009 | 1.7 Technical Requirements | 1.7.8 Constant and Consistent Conditions |  |  | V1M3 1.7.8 | Constant and Consistent Test Conditions Sample and Sampling Requirements
1010 | 1.7 Technical Requirements | 1.7.8 Constant and Consistent Conditions | Holding Time |  | V1M3 1.7.8.1 | Samples shall be transported to the laboratory as soon as possible after collection.
1011 | 1.7 Technical Requirements | 1.7.8 Constant and Consistent Conditions | Sampling Time and Date | Record | V1M3 1.7.8.1 | Date and time of sampling shall be noted on submittal forms.
1012 | 1.7 Technical Requirements | 1.7.8 Constant and Consistent Conditions | Collectors | Record | V1M3 1.7.8.1 | The names of the collectors with their signatures and the site shall be included on the chain-of-custody forms. No preservatives are required during sampling.
1013 | 1.7 Technical Requirements | 1.7.8 Constant and Consistent Conditions | Cross Contamination | Procedure | V1M3 1.7.8.2 | The laboratory shall establish and adhere to written procedures to minimize the possibility of cross contamination between samples.
1014 | 1.7 Technical Requirements | 1.7.8 Constant and Consistent Conditions | Method Requirements | Requirement | V1M3 1.7.8.3 | Refer to the specific method of analysis for additional requirements.

## Module 4
ID | Major Topic | Topic | Subject | Category | Citation | Does the laboratory comply with this section? | Yes | No | N/A
1015 |  |  |  |  | V1M4 1 | CHEMICAL TESTING
1016 | 1.1 Introduction |  |  |  | V1M4 1.1 | Introduction
1017 | 1.1 Introduction |  |  |  | V1M4 1.1 | This document contains detailed quality control requirements for environmental testing activities involving chemical measurements. The evaluation of laboratories for this discipline is in conjunction with a quality system as specified in the general requirements module. Adherence to quality systems requirements will ensure that all quality control procedures specified in this module are being followed.
1018 | 1.2 Scope |  |  |  | V1M4 1.2 | Scope
1019 | 1.2 Scope |  |  |  | V1M4 1.2 | The essential quality control procedures applicable to chemistry measurements are included in this module. Additional quality control requirements that are either specified by method, regulation or project shall be met by laboratories.
1020 | 1.3 Terms and Definitions |  |  |  | V1M4 1.3 | Terms and Definitions
1021 | 1.3 Terms and Definitions |  |  |  | V1M4 1.3 | The relevant definitions from TNI, Volume 1, Module 2, Section 3.0 are the preferred references. Definitions related to this document, which are used differently or do not exist in the above references are defined below.
1022 | 1.3 Terms and Definitions |  |  |  | V1M4 1.3.1 | Additional Terms and Definitions
1023 | 1.3 Terms and Definitions |  |  |  | V1M4 1.3.1 | Reserved
1024 | 1.3 Terms and Definitions |  |  |  | V1M4 1.3.2 | Exclusions and Exceptions
1025 | 1.3 Terms and Definitions |  |  |  | V1M4 1.3.2 | Reserved
1026 | 1.4 Method Selection |  |  |  | V1M4 1.4 | Method Selection
1027 | 1.4 Method Selection | Reference Method | Regulatory | Definition | V1M4 1.4 | A reference method is a method issued by an organization generally recognized as competent to do so. (When ISO refers to a standard method, that term is equivalent to reference method).
1028 | 1.4 Method Selection | Reference Method | Regulatory | Requirement | V1M4 1.4 | When a laboratory is required to analyze a parameter by a specified method due to a regulatory requirement, the parameter/method combination is recognized as a reference method.
1029 | 1.4 Method Selection | Reference Method | QC | Requirement | V1M4 1.4 | If there is not a regulatory requirement for the parameter/method combination, the parameter/ method combination need not be validated under 1.5.1b) as a non-reference method if it can be analyzed by another similar reference method of the same matrix and technology. The inclusion of the parameter in the method shall meet all required calibration requirements and the quality control requirements of the method to which the parameter is being added. If no QC exists in the method, the laboratory shall adhere to the requirements outlined in the similar method. For example, when adding acetone to Method 624, the calibration and QC requirements shall follow Method 624.
1030 | 1.4 Method Selection | Added Analytes | Identification | Requirement |  | A method that meets the above requirement shall be identified in such a way so that there is no confusion that the method has been modified.
1031 | 1.4 Method Selection | Non Reference Methods | Customer Communication | Requirement | V1M4 1.4 | When it is necessary to use methods not covered by reference methods, these shall be subject to agreement with the client and shall include a clear specification of the client's requirements and the purpose of the environmental test.
1032 | 1.4 Method Selection | Non Reference Methods | Validation | Requirement | V1M4 1.4 | The method developed shall have been validated appropriately before use.
1033 | 1.5 Method Validation |  |  |  | V1M4 1.5 | Method Validation
1034 | 1.5 Method Validation | 1.5.1 General |  |  | V1M4 1.5.1 | Validation of Methods
1035 | 1.5 Method Validation | 1.5.1 General | Reference Methods | Requirement | V1M4 1.5.1 a) | The laboratory shall validate reference methods via the procedures specified in Sections 1.5.2 and 1.5.3.
1036 | 1.5 Method Validation | 1.5.1 General | Non Reference Methods | Requirement | V1M4 1.5.1 b) | The laboratory shall validate non-reference methods, laboratory-designed/developed methods, reference methods used outside their published scope, and amplifications and modifications of reference methods to confirm that the methods are fit for the intended use. The validation shall be as extensive as is necessary to meet the needs of the given application or field of application. The laboratory shall record the results obtained, the procedure used for the validation, and a statement as to whether the method is fit for the intended use. In the absence of other specifications, the minimum requirements for method validation are given in Sections 1.5.2, 1.5.3 and 1.5.4.
1037 | 1.5 Method Validation | 1.5.2 LOD/LOQ |  |  | V1M4 1.5.2 | Limit of Detection and Limit of Quantitation (However Named)
1038 | 1.5 Method Validation | 1.5.2 LOD/LOQ | Documentation | Procedure | V1M4 1.5.2 | Procedures used for determining limits of detection and quantitation shall be documented. Documentation shall include the quality system matrix type. All supporting data shall be retained.
1039 | 1.5 Method Validation | 1.5.2.1 LOD |  |  | V1M4 1.5.2.1 | Limit of Detection (LOD)
1040 | 1.5 Method Validation | 1.5.2.1 LOD | LOD Study | Exception | V1M4 1.5.2.1 | If the laboratory is not reporting a value below the Limit of Quantitation, a Limit of Detection study is not required.
1041 | 1.5 Method Validation | 1.5.2.1 LOD | Appropriateness | Requirement | V1M4 1.5.2.1 | The laboratory shall utilize a method that provides an LOD that is appropriate and relevant for the intended use of the data.
1042 | 1.5 Method Validation | 1.5.2.1 LOD | Mandated Procedure | Requirement | V1M4 1.5.2.1 | If a mandated method or regulation includes protocols for determining detection limits, these shall be followed.
1043 | 1.5 Method Validation | 1.5.2.1 LOD | Selected Procedure | Record | V1M4 1.5.2.1 | The laboratory shall document how LODs were derived from the determinations. If the protocol for determining the LOD is not specified, the selection of the procedure shall reflect instrument limitations and the intended application of the method.
1044 | 1.5 Method Validation | 1.5.2.1 LOD | Processing | Requirement | V1M4 1.5.2.1 | All sample-processing and analysis steps of the analytical method shall be included in the determination or validation of the LOD.
1045 | 1.5 Method Validation | 1.5.2.1 LOD | Verification | Requirement | V1M4 1.5.2.1 a) | When required, the laboratory shall determine or verify the LOD for the method for each target analyte of concern in the quality system matrices.
1046 | 1.5 Method Validation | 1.5.2.1 LOD | Verification Procedure | Requirement | V1M4 1.5.2.1 b) | The validity of the LOD shall be verified by detection (a value above zero) of the analyte(s) in a QC sample in each quality system matrix. This QC sample shall contain the analyte at no more than 3X the LOD for single analyte tests and 4X the LOD for multiple analyte tests.
1047 | 1.5 Method Validation | 1.5.2.1 LOD | Extent of Verification | Requirement | V1M4 1.5.2.1 b) | This verification shall be performed on every instrument that is to be used for analysis of samples and reporting of data. The validity of the LOD shall be verified as part of the LOD determination process. This verification shall be done prior to the use of the LOD for the sample analysis.
1048 | 1.5 Method Validation | 1.5.2.1 LOD | LOD Study | Exception | V1M4 1.5.2.1 c) | An LOD study is not required for any component for which spiking solutions or quality control samples are not available such as temperature.
1049 | 1.5 Method Validation | 1.5.2.1 LOD | Initial Determination | Requirement | V1M4 1.5.2.1 d) | The LOD shall be initially determined for the compounds of interest in each method in a quality system matrix in which there are neither target analytes nor interferences at a concentration that would impact the results or the LOD shall be performed in the quality system matrix of interest.
1050 | 1.5 Method Validation | 1.5.2.1 LOD | Frequency | Requirement | V1M4 1.5.2.1 e) | An LOD shall be performed each time there is a change in the method that affects how the test is performed, or when a change in instrumentation occurs that affects the sensitivity of the analysis.
1051 | 1.5 Method Validation | 1.5.2.1 LOD | Frequency | Requirement | V1M4 1.5.2.1 f) | The LOD, if required, shall be verified annually for each quality system matrix, technology, and analyte.
1052 | 1.5 Method Validation | 1.5.2.2 LOQ |  |  | V1M4 1.5.2.2 | Limit of Quantitation (LOQ)
1053 | 1.5 Method Validation | 1.5.2.2 LOQ | Processing | Requirement | V1M4 1.5.2.2 a) | All sample-processing and analysis steps of the analytical method shall be included in the determination of the LOQ.
1054 | 1.5 Method Validation | 1.5.2.2 LOQ | LOQ Study | Exception | V1M4 1.5.2.2 b) | The LOQ study is not required for any component or property for which spiking solutions or quality control samples are not available or otherwise inappropriate (e.g., pH).
1055 | 1.5 Method Validation | 1.5.2.2 LOQ | Verification | Requirement | V1M4 1.5.2.2 c) | The validity of the LOQ shall be verified by successful analysis of a QC sample containing the analytes of concern in each quality system matrix at 1 to 2 times the claimed LOQ. A successful analysis is one where the recovery of each analyte is within the laboratory established method acceptance criteria or client data quality objectives for accuracy.
1056 | 1.5 Method Validation | 1.5.2.2 LOQ | Verification | Exception | V1M4 1.5.2.2 d) | When an LOD is determined or verified by the laboratory, the LOQ shall be above the LOD.
1057 | 1.5 Method Validation | 1.5.2.2 LOQ | Frequency | Requirement | V1M4 1.5.2.2 e) | The LOQ shall be verified annually for each quality system matrix, technology, and analyte. However, the annual LOQ verification is not required if the LOD was determined or verified annually on that instrument.
1058 | 1.5 Method Validation | 1.5.3 Precision & Bias |  |  | V1M4 1.5.3 | Evaluation of Precision and Bias
1059 | 1.5 Method Validation | 1.5.3 Precision & Bias | Reference Methods | Requirement | V1M4 1.5.3 a) | Reference Methods. The laboratory shall evaluate the precision and bias of a reference method for each analyte of concern for each quality system matrix according to Section 1.6 or alternate documented procedure when the analyte cannot be spiked into the sample matrix and QC samples are not commercially available.
1060 | 1.5 Method Validation | 1.5.3 Precision & Bias | Non Reference Methods | Requirement | V1M4 1.5.3 b) | Non-Reference Methods. For laboratory-developed methods or non-reference methods that were not in use by the laboratory before July 2003, the laboratory shall have a documented procedure to evaluate precision and bias. The laboratory shall also compare results of the precision and bias measurements with criteria established by the client, by criteria given in the reference method or criteria established by the laboratory.
1061 | 1.5 Method Validation | 1.5.3 Precision & Bias | Non Reference Methods | Requirement | V1M4 1.5.3 b) | Precision and bias measurements shall evaluate the method across the analytical calibration range of the method. The laboratory shall also evaluate precision and bias in the relevant quality system matrices and shall process the samples through the entire measurement system for each analyte of interest.
1062 | 1.5 Method Validation | 1.5.3 Precision & Bias | Non Reference Methods |  | V1M4 1.5.3 b) | Examples of a systematic approach to evaluate precision and bias could be the following:
1063 | 1.5 Method Validation | 1.5.3 Precision & Bias | Non Reference Methods | Option | V1M4 1.5.3 b) i. | Analyze QC samples in triplicate containing the analytes of concern at or near the limit of quantitation, at the upper-range of the calibration (upper 20%) and at a mid-range concentration. Process these samples on different days as three (3) sets of samples through the entire measurement system for each analyte of interest. Each day, one (1) QC sample at each concentration is analyzed. A separate method blank shall be subjected to the analytical method along with the QC samples on each of the three (3) days. (Note that the three (3) samples at the LOQ concentration can demonstrate sensitivity as well.) For each analyte, calculate the mean recovery for each day, for each level over each day, and for all nine (9) samples. Calculate the relative standard deviation for each of the separate means obtained. Compare the standard deviations for the different days and the standard deviations for the different concentrations. If the different standard deviations are all statistically insignificant (e.g., F-test), then compare the overall mean and standard deviation with the established criteria from above.
1064 | 1.5 Method Validation | 1.5.3 Precision & Bias | Non Reference Methods | Option | V1M4 1.5.3 b) ii. | A validation protocol, such as the Tier I, Tier II, and Tier III requirements in US EPA Office of Water’s Alternate Test Procedure (ATP) approval process.
1065 | 1.5 Method Validation | 1.5.4 Selectivity |  |  | V1M4 1.5.4 | Evaluation of Selectivity
1066 | 1.5 Method Validation | 1.5.4 Selectivity | Method Specified | Requirement | V1M4 1.5.4 | The laboratory shall evaluate selectivity by following the checks established within the method, which may include mass spectral tuning, second column confirmation, ICP inter-element interference checks, chromatography retention time windows, sample blanks, spectrochemical absorption or fluorescence profiles, co-precipitation evaluations, and electrode response factors.
1067 | 1.6 DOC |  |  |  | V1M4 1.6 | Demonstration of Capability (DOC)
1068 | 1.6 DOC | 1.6.1 General |  |  | V1M4 1.6.1 | General
1069 | 1.6 DOC | 1.6.1 General | Initial DOC | Requirement | V1M4 1.6.1 | Prior to acceptance and institution of any method for which data will be reported, a satisfactory initial DOC is required (see Section 1.6.2).
1070 | 1.6 DOC | 1.6.1 General | Ongoing DOC | Requirement | V1M4 1.6.1 | Thereafter, ongoing DOC (Section 1.6.3), as per the quality control requirements in Section 1.7.3 (such as laboratory control samples) is required.
1071 | 1.6 DOC | 1.6.1 General | DOC | Exception | V1M4 1.6.1 | In cases where a laboratory analyzes samples using a method that has been in use by the laboratory for at least one year prior to applying for accreditation, and there have been no significant changes in instrument type, personnel or method, the ongoing DOC shall be acceptable as an initial DOC.
1072 | 1.6 DOC | 1.6.1 General | DOC | Record | V1M4 1.6.1 | The laboratory shall have records on file to demonstrate that an initial DOC is not required.
1073 | 1.6 DOC | 1.6.1 General | Documentation | Record | V1M4 1.6.1 | For the initial DOC, appropriate records as discussed in Section 1.6.2 shall be completed.
1074 | 1.6 DOC | 1.6.1 General | Frequency | Requirement | V1M4 1.6.1 | An initial DOC shall be completed each time there is a change in instrument type, personnel, or method. (see 1.6.2)
1075 | 1.6 DOC | 1.6.1 General | Documentation | Requirement | V1M4 1.6.1 | All demonstrations shall be documented. .
1076 | 1.6 DOC | 1.6.1 General | Data | Record | V1M4 1.6.1 | All data applicable to the demonstration shall be retained and readily available at the laboratory.
1077 | 1.6 DOC | 1.6.2 Initial DOC |  |  | V1M4 1.6.2 | Initial DOC
1078 | 1.6 DOC | 1.6.2 Initial DOC | DOC Frequency | Requirement | V1M4 1.6.2 | An initial DOC shall be conducted prior to using any method, and at any time there is a change in instrument type, personnel or method or any time that a method has not been performed by the laboratory or analyst in a twelve (12) month period.
1079 | 1.6 DOC | 1.6.2 Initial DOC | Documentation | Record | V1M4 1.6.2.1 | The laboratory shall document each initial DOC in a manner such that the following information is readily available for each affected employee:
1080 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M4 1.6.2.1 a) | analyst(s) involved in preparation and/or analysis;
1081 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M4 1.6.2.1 b) | matrix;
1082 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M4 1.6.2.1 c) | analyte(s), class of analyte(s), or measured parameter(s);
1083 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M4 1.6.2.1 d) | identification of method(s) performed;
1084 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M4 1.6.2.1 e) | identification of laboratory-specific SOP used for analysis, including revision number;
1085 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M4 1.6.2.1 f) | date(s) of analysis; and
1086 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M4 1.6.2.1 g) | summary of analyses, including information outlined in Section 1.6.2.2.c.
1087 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M4 1.6.2.2 | If the method or regulation does not specify an initial DOC, the following procedure is acceptable. It is the responsibility of the laboratory to document that other approaches to initial DOC are adequate.
1088 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M4 1.6.2.2 a) | The analyte(s) shall be diluted in a volume of clean quality system matrix (a sample in which no target analytes or interferences are present at concentrations that will impact the results of a specific method) sufficient to prepare four (4) aliquots at the concentration specified, or if unspecified, to a concentration of one (1) to four (4) times the limit of quantitation.
1089 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M4 1.6.2.2 b) | At least four (4) aliquots shall be prepared and analyzed according to the method(s) either concurrently or over a period of days.
1090 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M4 1.6.2.2 c) | Using all of the results, calculate the mean recovery in the appropriate reporting units and the standard deviations of the sample (in the same units) for each parameter of interest. When it is not possible to determine mean and standard deviations, such as for presence/absence and logarithmic values, the laboratory shall assess performance against established and documented criteria.
1091 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M4 1.6.2.2 d) | Compare the information from (c) above to the corresponding acceptance criteria for precision and accuracy in the method (if applicable) or in laboratory-generated acceptance criteria (if there are not established mandatory criteria). If all parameters meet the acceptance criteria, the analysis of actual samples may begin. If any one of the parameters does not meet the acceptance criteria, the performance is unacceptable for that parameter.
1092 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M4 1.6.2.2 e) | When one or more of the tested parameters fail at least one of the acceptance criteria, the analyst shall proceed according to i) or ii) below.
1093 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M4 1.6.2.2 e) i. | Locate and correct the source of the problem and repeat the test for all parameters of interest beginning with b) above.
1094 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M4 1.6.2.2 e) ii. | Beginning with b) above, repeat the test for all parameters that failed to meet criteria.
1095 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M4 1.6.2.2 f) | Repeated failure, however, confirms a general problem with the measurement system. If this occurs, locate and correct the source of the problem and repeat the test for all compounds of interest beginning with b).
1096 | 1.6 DOC | 1.6.2 Initial DOC | Added Analyte | Requirement | V1M4 1.6.2.2 g) | When an analyte not currently found on the laboratory’s list of accredited analytes is added to an existing accredited method, an initial demonstration shall be performed for that analyte.
1097 | 1.6 DOC | 1.6.3 Ongoing DOC |  |  | V1M4 1.6.3 | Ongoing DOC
1098 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Requirement | V1M4 1.6.3.1 | The laboratory shall have a documented procedure describing ongoing DOC. The analyst(s) shall demonstrate on-going capability by meeting the quality control requirements of the method, laboratory SOP, client specifications, and/or this Standard. It is the responsibility of the laboratory to document that other approaches to ongoing DOC are adequate.
1099 | 1.6 DOC | 1.6.3 Ongoing DOC |  |  | V1M4 1.6.3.2 | This on-going demonstration may be one of the following:
1100 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M4 1.6.3.2 a) | acceptable performance of a blind sample (single blind to the analyst);
1101 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M4 1.6.3.2 a) | Note: Successful analysis of a blind performance sample on a similar method using the same technology (e.g., GC/MS volatiles by purge and trap for Methods 524.2, 624 or 5030/8260) would only require documentation for one of the tests.
1102 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M4 1.6.3.2 b) | another initial DOC;
1103 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M4 1.6.3.2 c) | at least four (4) consecutive laboratory control samples with acceptable levels of precision and accuracy. The laboratory shall determine the acceptable limits for precision and accuracy prior to analysis. The laboratory shall tabulate or be able to readily retrieve four (4) consecutive passing LCSs for each method for each analyst each year;
1104 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M4 1.6.3.2 d) | a documented process of analyst review using QC samples. QC samples can be reviewed to identify patterns for individuals or groups of analysts and determine if corrective action or retraining is necessary;
1105 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M4 1.6.3.2 e) | if a) through d) are not technically feasible, then analysis of real-world samples with results within a predefined acceptance criteria (as defined by the laboratory or method) shall be performed.
1106 | 1.7 Technical Requirement |  |  |  | V1M4 1.7 | Technical Requirements
1107 | 1.7 Technical Requirement | 1.7.1.Initial Calibration |  |  | V1M4 1.7.1 | Initial Calibration
1108 | 1.7 Technical Requirement | 1.7.1.Initial Calibration |  |  | V1M4 1.7.1.1 | Instrument Calibration
1109 | 1.7 Technical Requirement | 1.7.1.Initial Calibration |  |  | V1M4 1.7.1.1 | This module specifies the essential elements that shall define the procedures and documentation for initial instrument calibration and continuing instrument calibration verification to ensure that the data shall be of known quality for the intended use. This Standard does not specify detailed procedural steps (“how to”) for calibration, but establishes the essential elements for selection of the appropriate technique(s). This approach allows flexibility and permits the employment of a wide variety of analytical procedures and statistical approaches currently applicable for calibration.
1110 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Multiple/Conflicting Requirements | Requirement | V1M4 1.7.1.1 | If more stringent standards or requirements are included in a mandated method or by regulation, the laboratory shall demonstrate that such requirements are met. If it is not apparent which Standard is more stringent, then the requirements of the regulation or mandated method are to be followed.
1111 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Essential Elements |  | V1M4 1.7.1.1 | The following items are essential elements of initial instrument calibration:
1112 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Process | Procedure | V1M4 1.7.1.1 a) | the details of the initial instrument calibration procedures including calculations, integrations, acceptance criteria and associated statistics shall be included or referenced in the method SOP. When initial instrument calibration procedures are referenced in the method, then the referenced material shall be retained by the laboratory and be available for review;
1113 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Data | Record | V1M4 1.7.1.1 b) | sufficient raw data records shall be retained to permit reconstruction of the initial instrument calibration (e.g., calibration date, method, instrument, analysis date, each analyte name, analyst’s initials or signature; concentration and response, calibration curve or response factor; or unique equation or coefficient used to reduce instrument responses to concentration);
1114 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Quantitation | Requirement | V1M4 1.7.1.1 c) | sample results shall be quantitated from the initial instrument calibration and may not be quantitated from any continuing instrument calibration verification unless otherwise required by regulation, method, or program;
1115 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Second Source Standard | Requirement | V1M4 1.7.1.1 d) | all initial instrument calibrations shall be verified with a standard obtained from a second manufacturer or from a different lot. Traceability shall be to a national standard, when commercially available;
1116 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Acceptance Criteria | Requirement | V1M4 1.7.1.1 e) | criteria for the acceptance of an initial instrument calibration shall be established (e.g., correlation coefficient or relative percent difference). The criteria used shall be appropriate to the calibration technique employed;
1117 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Lowest Calibration Concentration | Requirement | V1M4 1.7.1.1 f) | the lowest calibration standard shall be at or below the LOQ. Any data reported below the LOQ shall be considered to have an increased quantitative uncertainty and shall be reported using defined qualifiers or explained in the narrative;
1118 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Highest Calibration Concentration | Requirement | V1M4 1.7.1.1 g) | the highest calibration standard shall be at or above the highest concentration for which quantitative data are to be reported. Any data reported above the calibration range shall be considered to have an increased quantitative uncertainty and shall be reported using defined qualifiers or explained in the narrative;
1119 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | One-Point Calibration | Requirement | V1M4 1.7.1.1 h) | the following shall occur for instrument technology (such as ICP or ICP/MS) with validated techniques from manufacturers or methods employing standardization with a zero point and a single point calibration standard:
1120 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | One-Point Calibration | Requirement | V1M4 1.7.1.1 h) i. | Prior to the analysis of samples, the zero point and single point calibration standard shall be analyzed and the linear range of the instrument shall be established by analyzing a series of standards, one of which shall be at or below the LOQ. Sample results within the established linear range will not require data qualifiers.
1121 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | One-Point Calibration | Requirement | V1M4 1.7.1.1 h) ii. | A zero point and single point calibration standard shall be analyzed with each analytical batch.
1122 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | One-Point Calibration | Requirement | V1M4 1.7.1.1 h) iii. | A standard corresponding to the limit of quantitation shall be analyzed with each analytical batch and shall meet established acceptance criteria.
1123 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | One-Point Calibration | Requirement | V1M4 1.7.1.1 h) iv. | The linearity is verified at a frequency established by the method and/or the manufacturer.
1124 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Corrective Actions | Requirement | V1M4 1.7.1.1 i) | if the initial instrument calibration results are outside established acceptance criteria, corrective actions shall be performed and all associated samples re-analyzed. If re-analysis of the samples is not possible, data associated with an unacceptable initial instrument calibration shall be reported with appropriate data qualifiers; and
1125 | 1.7 Technical Requirement | 1.7.1.Initial Calibration | Minimum Calibration Points | Requirement | V1M4 1.7.1.1 j) | if a reference or mandated method does not specify the number of calibration standards, the minimum number of points for establishing the initial instrument calibration shall be three.
1126 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration |  |  | V1M4 1.7.2 | Continuing Calibration
1127 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Frequency | Requirement | V1M4 1.7.2 | When an initial instrument calibration is not performed on the day of analysis, the validity of the initial calibration shall be verified prior to sample analyses by a continuing instrument calibration verification with each analytical batch. The following items are essential elements of continuing instrument calibration verification.
1128 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Documentation | Procedure | V1M4 1.7.2 a) | The details of the continuing instrument calibration procedure, calculations and associated statistics shall be included or referenced in the method SOP.
1129 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Extent of Verification | Requirement | V1M4 1.7.2 b) | Calibration shall be verified for each compound, element, or other discrete chemical species, except for multi-component analytes such as aroclors, chlordane, total petroleum hydrocarbons, or toxaphene, where a representative chemical, related substance or mixture can be used.
1130 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | CC Performance |  | V1M4 1.7.2 c) | Instrument calibration verification shall be performed:
1131 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Batch Frequency | Requirement | V1M4 1.7.2 c) i. | at the beginning and end of each analytical batch. If an internal standard is used, only one verification needs to be performed at the beginning of the analytical batch;
1132 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Time Frequency | Requirement | V1M4 1.7.2 c) ii. | if the time period for calibration or the most recent calibration verification has expired; or
1133 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Method Frequency | Requirement | V1M4 1.7.2 c) iii. | for analytical systems that contain a calibration verification requirement.
1134 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Calibration Reconstruction | Record | V1M4 1.7.2 d) | Sufficient raw data records shall be retained to permit reconstruction of the continuing instrument calibration verification (e.g., method, instrument, analysis date, each analyte name, concentration and response, calibration curve or response factor, or unique equations or coefficients used to convert instrument responses into concentrations). Continuing calibration verification records shall explicitly connect the continuing verification data to the initial instrument calibration.
1135 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Acceptance Criteria | Requirement | V1M4 1.7.2 e) | Criteria for the acceptance of a continuing instrument calibration verification shall be established.
1136 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Corrective Actions | Requirement | V1M4 1.7.2 e) | If the continuing instrument calibration verification results obtained are outside the established acceptance criteria and analysis of a second consecutive (immediate) calibration verification fails to produce results within acceptance criteria, corrective actions shall be performed.
1137 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Verification of Correction Action | Requirement | V1M4 1.7.2 e) | The laboratory shall demonstrate acceptable performance after corrective action with two consecutive calibration verifications, or a new initial instrument calibration shall be performed.
1138 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Process without verified calibration | Requirement | V1M4 1.7.2 e) | If the laboratory has not verified calibration, sample analyses may not occur until the analytical system is calibrated or calibration verified.
1139 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Process without verified calibration | Requirement | V1M4 1.7.2 e) | If samples are analyzed using a system on which the calibration has not yet been verified the results shall be flagged.
1140 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Process without verified calibration | Requirement | V1M4 1.7.2 e) | Data associated with an unacceptable calibration verification may be fully useable under the following special conditions:
1141 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Data Usability | Option | V1M4 1.7.2 e) i. | when the acceptance criteria for the continuing calibration verification are exceeded high (i.e., high bias) and there are associated samples that are non-detects, then those non-detects may be reported. Otherwise the samples affected by the unacceptable calibration verification shall be re-analyzed after a new calibration curve has been established, evaluated and accepted; or
1142 | 1.7 Technical Requirement | 1.7.2 Continuing Calibration | Data Usability | Option | V1M4 1.7.2 e) ii. | when the acceptance criteria for the continuing calibration verification are exceeded low (i.e., low bias), those sample results may be reported if they exceed a maximum regulatory limit/decision level. Otherwise the samples affected by the unacceptable verification shall be re-analyzed after a new calibration curve has been established, evaluated and accepted.
1143 | 1.7 Technical Requirement | 1.7.3 QC |  |  | V1M4 1.7.3 | Quality Control
1144 | 1.7 Technical Requirement |  |  |  | V1M4 1.7.3 | The laboratory shall have quality control procedures for monitoring the validity of environmental tests undertaken as specified in this Section.(see requirements below)
1145 | 1.7 Technical Requirement | 1.7.3.1 Negative Control |  |  | V1M4 1.7.3.1 | Negative Control – Method Performance: Method Blank
1146 | 1.7 Technical Requirement | 1.7.3.1 Negative Control | Method Blank Processing | Requirement | V1M4 1.7.3.1 a) | The method blank is used to assess the samples in the preparation batch for possible contamination during the preparation and processing steps.
1147 | 1.7 Technical Requirement | 1.7.3.1 Negative Control | Method Blank Processing | Requirement | V1M4 1.7.3.1 a) | The method blank shall be processed along with and under the same conditions as the associated samples to include all steps of the analytical procedure.
1148 | 1.7 Technical Requirement | 1.7.3.1 Negative Control | Blank Assessment | Procedure | V1M4 1.7.3.1 a) | Procedures shall be in place to determine if a method blank is contaminated.
1149 | 1.7 Technical Requirement | 1.7.3.1 Negative Control | Reporting | Requirement | V1M4 1.7.3.1 a) | Any affected samples associated with a contaminated method blank shall be reprocessed for analysis or the results reported with appropriate data qualifying codes.
1150 | 1.7 Technical Requirement | 1.7.3.1 Negative Control | Frequency | Requirement | V1M4 1.7.3.1 b) | The method blank shall be analyzed at a minimum of one (1) per preparation batch.
1151 | 1.7 Technical Requirement | 1.7.3.1 Negative Control | Frequency | Requirement | V1M4 1.7.3.1 b) | In those instances for which no separate preparation method is used (for example, volatiles in water), the batch shall be defined as environmental samples that are analyzed together with the same method and personnel, using the same lots of reagents, not to exceed the analysis of twenty (20) environmental samples, not including method blanks, LCS, matrix spikes and matrix duplicates.
1152 | 1.7 Technical Requirement | 1.7.3.1 Negative Control | Matrix Attributes | Requirement | V1M4 1.7.3.1 c) | The method blank shall consist of a quality system matrix that is similar to the associated samples and is known to be free of the analytes of interest.
1153 | 1.7 Technical Requirement | 1.7.3.1 Negative Control | Performing | Exception | V1M4 1.7.3.1 d) | Method blanks are not applicable for certain analyses, such as pH, Conductivity, Flash Point and Temperature.
1154 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls |  |  | V1M4 1.7.3.2 | Positive Control – Method Performance: Laboratory Control Sample (LCS)
1155 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Purpose and Use | Requirement | V1M4 1.7.3.2.1 | The LCS is used to evaluate the performance of the total analytical system, including all preparation and analysis steps. Results of the LCS are compared to established criteria and, if found to be outside of these criteria, indicates that the analytical system is “out of control.” Any affected samples associated with an out of control LCS shall be reprocessed for re-analysis or the results reported with appropriate data qualifying codes.
1156 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Frequency | Requirement | V1M4 1.7.3.2.2 | The LCS shall be analyzed at a minimum of one (1) per preparation batch.
1157 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Performance | Exception | V1M4 1.7.3.2.2 | Exceptions would be for those analytes for which no spiking solutions are available, such as total suspended solids, total dissolved solids, total volatile solids, total solids, pH, color, odor, temperature, dissolved oxygen or turbidity.
1158 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Frequency | Requirement | V1M4 1.7.3.2.2 | In those instances for which no separate preparation method is used (example: volatiles in water) the batch shall be defined as environmental samples that are analyzed together with the same method and personnel, using the same lots of reagents, not to exceed the analysis of twenty (20) environmental samples, not including method blanks, LCS, matrix spikes and matrix duplicates.
1159 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Matrix Attributes | Requirement | V1M4 1.7.3.2.3 | The LCS is a quality system matrix, known to be free of analytes of interest, spiked with known concentrations of analytes.
1160 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Substitution | Option | V1M4 1.7.3.2.3 | Note: The matrix spike may be used in place of this control as long as the acceptance criteria are as stringent as for the LCS.
1161 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | CRM | Option | V1M4 1.7.3.2.3 | Alternatively, the LCS may consist of a media containing known and verified concentrations of analytes or as Certified Reference Material (CRM). All analyte concentrations shall be within the calibration range of the methods. The following shall be used in choosing components for the spike mixtures:
1162 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Spike Components | Requirement | V1M4 1.7.3.2.3 | The components to be spiked shall be as specified by the mandated method or regulation or as requested by the client. In the absence of specified spiking components, the laboratory shall spike per the following:
1163 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Spike Components | Option | V1M4 1.7.3.2.3 a) | for those components that interfere with an accurate assessment, such as spiking simultaneously with technical chlordane, toxaphene and PCBs, the spike shall be chosen that represents the chemistries and elution patterns of the components to be reported; and
1164 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Spike Components | Option | V1M4 1.7.3.2.3 b) | for those methods that have extremely long lists of analytes, a representative number may be chosen. The analytes selected shall be representative of all analytes reported.
1165 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Multi Analytes Methods | Requirement | V1M4 1.7.3.2.3 b) | The following criteria shall be used for determining the minimum number of analytes to be spiked. However, the laboratory shall insure that all targeted components are included in the spike mixture over a two (2) year period:
1166 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Multi Analytes Methods | Option | V1M4 1.7.3.2.3 b) i. | For methods that include one (1) to ten (10) targets, spike all components.
1167 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Multi Analytes Methods | Option | V1M4 1.7.3.2.3 b) ii. | For methods that include eleven (11) to twenty (20) targets, spike at least ten (10) or 80%, whichever is greater.
1168 | 1.7 Technical Requirement | 1.7.3.2 Positive Controls | Multi Analytes Methods | Option | V1M4 1.7.3.2.3 b) iii. | For methods with more than twenty (20) targets, spike at least sixteen (16) components.
1169 | 1.7 Technical Requirement | 1.7.3.3 Sample Specific Controls |  |  | V1M4 1.7.3.3 | Sample Specific Controls
1170 | 1.7 Technical Requirement | 1.7.3.3 Sample Specific Controls | Documentation | Procedure | V1M4 1.7.3.3 | The laboratory shall document procedures for determining the effect of the sample matrix on method performance. These procedures relate to the analyses of quality system matrix specific Quality Control (QC) samples and are designed as data quality indicators for a specific sample using the designated method. These controls alone are not used to judge laboratory performance.
1171 | 1.7 Technical Requirement | 1.7.3.3 Sample Specific Controls | Documentation | Procedure | V1M4 1.7.3.3 | Examples of matrix-specific QC include: Matrix Spike (MS), Matrix Spike Duplicate (MSD), sample duplicates, and surrogate spikes. The laboratory shall have procedures in place for tracking, managing, and handling matrix-specific QC criteria, including spiking appropriate components at appropriate concentrations, calculating recoveries and relative percent difference, and evaluating and reporting results based on performance of the QC samples.
1172 | 1.7 Technical Requirement | 1.7.3.3.1 MS/MSD | MS/MSD |  | V1M4 1.7.3.3.1 | Matrix Spike; Matrix Spike Duplicates
1173 | 1.7 Technical Requirement | 1.7.3.3.1 MS/MSD | Use | Requirement | V1M4 1.7.3.3.1 a) | Matrix-specific QC samples indicate the effect of the sample matrix on the precision and accuracy of the results generated using the selected method. The information from these controls is sample/matrix specific and would not normally be used to determine the validity of the entire batch.
1174 | 1.7 Technical Requirement | 1.7.3.3.1 MS/MSD | Frequency | Requirement | V1M4 1.7.3.3.1 b) | The frequency of the analysis of matrix spikes are as specified by the method or may be determined as part of the contract review process.
1175 | 1.7 Technical Requirement | 1.7.3.3.1 MS/MSD | Spike Components | Requirement | V1M4 1.7.3.3.1 c) | The components to be spiked shall be as specified by the mandated method. Any permit specified analytes, as specified by regulation or client requested analytes, shall also be included. If there are no specified components, the laboratory shall spike per the following:
1176 | 1.7 Technical Requirement | 1.7.3.3.1 MS/MSD | Multi-peak Components | Requirement | V1M4 1.7.3.3.1 c) | For those components that interfere with an accurate assessment such as spiking simultaneously with technical chlordane, toxaphene and PCBs, the spike shall be chosen that represents the chemistries and elution patterns of the components to be reported.
1177 | 1.7 Technical Requirement | 1.7.3.3.1 MS/MSD | Spike Components | Requirement | V1M4 1.7.3.3.1 c) | For those methods that have extremely long lists of analytes, a representative number may be chosen using the following criteria for choosing the number of analytes to be spiked. However, the laboratory shall insure that all targeted components are included in the spike mixture over a two (2) year period.
1178 | 1.7 Technical Requirement | 1.7.3.3.1 MS/MSD | Spike Components | Option | V1M4 1.7.3.3.1 c) i. | For methods that include one (1) to ten (10) targets, spike all components.
1179 | 1.7 Technical Requirement | 1.7.3.3.1 MS/MSD | Spike Components | Option | V1M4 1.7.3.3.1 c) ii. | For methods that include eleven (11) to twenty (20) targets, spike at least ten (10) or 80%, whichever is greater.
1180 | 1.7 Technical Requirement | 1.7.3.3.1 MS/MSD | Spike Components | Option | V1M4 1.7.3.3.1 c) iii. | For methods with more than twenty (20) targets, spike at least sixteen (16) components.
1181 | 1.7 Technical Requirement | 1.7.3.3.2 Matrix Duplicates |  |  | V1M4 1.7.3.3.2 | Matrix Duplicates
1182 | 1.7 Technical Requirement | 1.7.3.3.2 Matrix Duplicates | Use |  | V1M4 1.7.3.3.2 a) | Matrix duplicates are defined as replicate aliquots of the same sample taken through the entire analytical procedure. The results from this analysis indicate the precision of the results for the specific sample using the selected method. The matrix duplicate may provide a usable measure of sample homogeneity. It may also provide a measure of precision when target analytes are present.
1183 | 1.7 Technical Requirement | 1.7.3.3.2 Matrix Duplicates | Frequency | Requirement | V1M4 1.7.3.3.2 b) | The frequency of the analysis of matrix duplicates are as specified by the method or may be determined as part of the contract review process.
1184 | 1.7 Technical Requirement | 1.7.3.3.2 Matrix Duplicates | Origin | Requirement | V1M4 1.7.3.3.2 c) | Matrix duplicates are performed on replicate aliquots of actual samples. The composition is usually not known.
1185 | 1.7 Technical Requirement | 1.7.3.3.3 Surrogate Spikes |  |  | V1M4 1.7.3.3.3 | Surrogate Spikes
1186 | 1.7 Technical Requirement | 1.7.3.3.3 Surrogate Spikes | Preparation | Requirement | V1M4 1.7.3.3.3 a) | Surrogates, when required, are chosen to reflect the chemistries of the targeted components of the method and are added prior to sample preparation/extraction.
1187 | 1.7 Technical Requirement | 1.7.3.3.3 Surrogate Spikes | Extent | Requirement | V1M4 1.7.3.3.3 b) | Except where the matrix precludes its use or when not commercially available, surrogate compounds shall be added to all samples, standards, and blanks for all appropriate methods.
1188 | 1.7 Technical Requirement | 1.7.3.3.3 Surrogate Spikes | Source | Requirement | V1M4 1.7.3.3.3 c) | Surrogate compounds are chosen to represent the various chemistries of the target analytes in the method. They are often specified by the mandated method and are deliberately chosen for their being unlikely to occur as an environmental contaminant. Often this is accomplished by using deuterated analogs of select compounds.
1189 | 1.7 Technical Requirement | 1.7.3.4 Data Reduction |  |  | V1M4 1.7.3.4 | Data Reduction
1190 | 1.7 Technical Requirement | 1.7.3.4 Data Reduction | Documentation | Procedure | V1M4 1.7.3.4 | The procedures for data reduction, such as use of linear regression, shall be documented.
1191 | 1.7 Technical Requirement | 1.7.3.5 Reagent & Water Quality |  |  | V1M4 1.7.3.5 | Reagent Quality, Water Quality and Checks
1192 | 1.7 Technical Requirement | 1.7.3.5 Reagent & Water Quality | Grade | Requirement | V1M4 1.7.3.5 a) | In methods where the purity of reagents is not specified, analytical reagent grade shall be used. Reagents of lesser purity than those specified by the method shall not be used. Documentation of purity shall be available.
1193 | 1.7 Technical Requirement | 1.7.3.5 Reagent & Water Quality | Water Monitoring | Requirement | V1M4 1.7.3.5 b) | The quality of water sources shall be monitored and documented and shall meet method specified requirements.
1194 | 1.7 Technical Requirement | 1.7.3.5 Reagent & Water Quality | Titrant Verification | Requirement | V1M4 1.7.3.5 c) | The laboratory shall verify the concentration of titrants in accordance with written laboratory procedures.
1195 | 1.7 Technical Requirement | 1.7.3.6 Selectivity |  |  | V1M4 1.7.3.6 | Selectivity
1196 | 1.7 Technical Requirement | 1.7.3.6 Selectivity | Documentation | Procedure | V1M4 1.7.3.6 | The laboratory shall document selectivity by following the checks established within the method.
1197 | 1.7 Technical Requirement | 1.7.4 Acceptance/Rejection Criteria |  |  | V1M4 1.7.4 | Data Acceptance/Rejection Criteria
1198 | 1.7 Technical Requirement | 1.7.4.1 Negative Control |  |  | V1M4 1.7.4.1 | Negative Control – Method Performance: Method Blank
1199 | 1.7 Technical Requirement | 1.7.4.1 Negative Control | Evaluation | Requirement | V1M4 1.7.4.1 | While the goal is to have no detectable contaminants, each method blank shall be critically evaluated as to the nature of the interference and the effect on the analysis of each sample within the batch.
1200 | 1.7 Technical Requirement | 1.7.4.1 Negative Control | Investigation | Requirement | V1M4 1.7.4.1 | The source of contamination shall be investigated and measures taken to minimize or eliminate the problem and affected samples reprocessed or data shall be appropriately qualified if: (1.7.4.1.c)
1201 | 1.7 Technical Requirement | 1.7.4.1 Negative Control | Data Qualifiers | Requirement | V1M4 1.7.4.1 a) | the concentration of a targeted analyte in the blank is at or above the reporting limit as established by the method or by regulation, AND is greater than 1/10 of the amount measured in the sample;
1202 | 1.7 Technical Requirement | 1.7.4.1 Negative Control | Data Qualifiers | Requirement | V1M4 1.7.4.1 b) | the blank contamination otherwise affects the sample results as per the method requirements or the individual project data quality objectives; and a blank is determined to be contaminated.
1203 | 1.7 Technical Requirement | 1.7.4.1 Negative Control | Investigation | Requirement | V1M4 1.7.4.1 c) | The cause shall be investigated and measures taken to minimize or eliminate the problem. Samples associated with a contaminated blank shall be evaluated as to the best corrective action for the samples (e.g., reprocessing or data qualifying codes). In all cases the corrective action shall be documented.
1204 | 1.7 Technical Requirement | 1.7.4.2 Positive Control |  |  | V1M4 1.7.4.2 | Positive Control – Method Performance: Laboratory Control Sample (LCS)
1205 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Calculation | Requirement | V1M4 1.7.4.2 a) | The results of the individual batch LCS are calculated in percent recovery or other appropriate statistical technique that allows comparison to established acceptance criteria. The laboratory shall document the calculation.
1206 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Evaluation | Requirement | V1M4 1.7.4.2 a) | The individual LCS is compared to the acceptance criteria as published in the mandated method. Where there are no established criteria, the laboratory shall determine internal criteria and document the method used to establish the limits or utilize client specified assessment criteria.
1207 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Unacceptable Results | Requirement | V1M4 1.7.4.2 a) | An LCS that is determined to be within the criteria effectively establishes that the analytical system is in control and validates system performance for the samples in the associated batch. Samples analyzed along with an LCS determined to be “out of control” shall be considered suspect and the samples reprocessed and re-analyzed or the data reported with appropriate data qualifying codes. This includes any allowable marginal exceedance as described in b) below.
1208 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Unacceptable Results | Option | V1M4 1.7.4.2 a) i. | when the acceptance criteria for the positive control are exceeded high (i.e., high bias) and there are associated samples that are non-detects, then those non-detects may be reported with data qualifying codes; or
1209 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Unacceptable Results | Option | V1M4 1.7.4.2 a) ii. | when the acceptance criteria for the positive control are exceeded low (i.e., low bias), those sample results may be reported if they exceed a maximum regulatory limit/decision level with data qualifying codes.
1210 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Allowable Marginal Exceedances | Requirement | V1M4 1.7.4.2 b) | Allowable Marginal Exceedances. If a large number of analytes are in the LCS, it becomes statistically likely that a few will be outside control limits. This may not indicate that the system is out of control, therefore corrective action may not be necessary. Upper and lower marginal exceedance (ME) limits can be established to determine when corrective action is necessary. A ME is defined as being beyond the LCS control limit (three standard deviations), but within the ME limits. ME limits are between three (3) and four (4) standard deviations around the mean. The number of allowable marginal exceedances is based on the number of analytes in the LCS. If more analytes exceed the LCS control limits than is allowed, or if any one analyte exceeds the ME limits, the LCS fails and corrective action is necessary. This marginal exceedance approach is relevant for methods with long lists of analytes. It will not apply to target analyte lists with fewer than eleven analytes.
1211 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Allowable Marginal Exceedances |  | V1M4 1.7.4.2 b) | The number of allowable marginal exceedances is as follows:
1212 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Allowable Marginal Exceedances | Option | V1M4 1.7.4.2 b) | Number of Analytes in LCS >90 Number Allowed as Marginal Exceedances 5
1213 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Allowable Marginal Exceedances | Option | V1M4 1.7.4.2 b) | Number of Analytes in LCS 71-90 Number Allowed as Marginal Exceedances 4
1214 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Allowable Marginal Exceedances | Option | V1M4 1.7.4.2 b) | Number of Analytes in LCS 51-70 Number Allowed as Marginal Exceedances 3
1215 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Allowable Marginal Exceedances | Option | V1M4 1.7.4.2 b) | Number of Analytes in LCS 31-50 Number Allowed as Marginal Exceedances 2
1216 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Allowable Marginal Exceedances | Option | V1M4 1.7.4.2 b) | Number of Analytes in LCS 11-30 Number Allowed as Marginal Exceedances 1
1217 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Allowable Marginal Exceedances | Option | V1M4 1.7.4.2 b) | Number of Analytes in LCS <11 Number Allowed as Marginal Exceedances 0
1218 | 1.7 Technical Requirement | 1.7.4.2 Positive Control | Corrective Actions | Requirement | V1M4 1.7.4.2 b) | If the same analyte exceeds the LCS control limit consecutively, it is an indication of a systemic problem. The source of the error shall be located and corrective action taken. Laboratories shall have a written procedure to monitor the application of marginal exceedance allowance to the LCS.
1219 | 1.7 Technical Requirement | 1.7.4.3 Sample Specific Controls |  |  | V1M4 1.7.4.3 | Sample Specific Controls
1220 | 1.7 Technical Requirement | 1.7.4.3a MS/MSD |  |  | V1M4 1.7.4.3 a) | Matrix Spike; Matrix Spike Duplicates
1221 | 1.7 Technical Requirement | 1.7.4.3a MS/MSD | Documentation | Procedure | V1M4 1.7.4.3 a) | The results from matrix spike/matrix spike duplicate are primarily designed to assess the precision and accuracy of analytical results in a given matrix and are expressed as percent recovery (%R), relative percent difference (RPD), or other appropriate statistical technique that allows comparison to established acceptance criteria. The laboratory shall document the calculation for %R, RPD or other statistical treatment used.
1222 | 1.7 Technical Requirement | 1.7.4.3a MS/MSD | Evaluation | Requirement | V1M4 1.7.4.3 a) | The results are compared to the acceptance criteria as published in the mandated method. Where there are no established criteria, the laboratory shall determine internal criteria and document the method used to establish the limits. For matrix spike results outside established criteria, corrective action shall be documented or the data for that sample reported with appropriate data qualifying codes.
1223 | 1.7 Technical Requirement | 1.7.4.3.b Matrix Duplicates |  |  | V1M4 1.7.4.3 b) | Matrix Duplicates
1224 | 1.7 Technical Requirement | 1.7.4.3.b Matrix Duplicates |  |  | V1M4 1.7.4.3 b) | The results from matrix duplicates are primarily designed to assess the homogeneity of the particular sample chosen. If that sample is homogenous it may also describe the precision of analytical results in a given matrix.
1225 | 1.7 Technical Requirement | 1.7.4.3.b Matrix Duplicates | Documentation | Procedure | V1M4 1.7.4.3 b) | These may be expressed as relative percent difference (RPD) or another statistical treatment (e.g., absolute differences). The laboratory shall document the calculation for relative percent difference or other statistical treatments.
1226 | 1.7 Technical Requirement | 1.7.4.3.b Matrix Duplicates | Evaluation | Requirement | V1M4 1.7.4.3 b) | Results are compared to the acceptance criteria as published in the mandated method. Where there are no established criteria, the laboratory shall determine internal criteria and document the method used to establish the limits. For matrix duplicates results outside established criteria, corrective action shall be documented or the data for that sample reported with appropriate data qualifying codes.
1227 | 1.7 Technical Requirement | 1.7.4.3.c Surrogates |  |  | V1M4 1.7.4.3 c) | Surrogate Spikes
1228 | 1.7 Technical Requirement | 1.7.4.3.c Surrogates | Acceptance Criteria | Requirement | V1M4 1.7.4.3 c) | The results are compared to the acceptance criteria as published in the mandated method. Where there are no established criteria, the laboratory shall determine internal criteria and document the method used to establish the limits.
1229 | 1.7 Technical Requirement | 1.7.4.3.c Surrogates | Evaluation | Requirement | V1M4 1.7.4.3 c) | Surrogates outside the acceptance criteria shall be evaluated for the effect indicated for the individual sample results. The appropriate corrective action may be guided by the data quality objectives or other site-specific requirements.
1230 | 1.7 Technical Requirement | 1.7.4.3.c Surrogates | Reporting | Requirement | V1M4 1.7.4.3 c) | Results reported from analyses with surrogate recoveries outside the acceptance criteria shall include appropriate data qualifiers.
1231 | 1.7 Technical Requirement | 1.7.5 Sample handing |  |  | V1M4 1.7.5 | Sample Handling
1232 | 1.7 Technical Requirement | 1.7.5 Sample handing | Temperature | Requirement | V1M4 1.7.5 a) | All samples that require thermal preservation shall be considered acceptable if the arrival temperature of a representative sample container is either within 2°C of the required temperature or the method specified range. For samples with a specified temperature of 4°C, samples with a temperature ranging from just above the freezing temperature of water to 6°C shall be acceptable.
1233 | 1.7 Technical Requirement | 1.7.5 Sample handing | Temperature | Exception | V1M4 1.7.5 a) i. | Samples that are delivered to the laboratory on the same day they are collected may not meet the requirements of Section 1.7.5.a. In these cases, the samples shall be considered acceptable if the samples were received on ice.
1234 | 1.7 Technical Requirement | 1.7.5 Sample handing | Temperature | Exception | V1M4 1.7.5 a) ii. | If sample analysis is begun within fifteen (15) minutes of collection, thermal preservation is not required.
1235 | 1.7 Technical Requirement | 1.7.5 Sample handing | Temperature | Exception | V1M4 1.7.5 a) iii. | Thermal preservation is not required in the field if the laboratory receives and refrigerates the sample within fifteen (15) minutes of collection.
1236 | 1.7 Technical Requirement | 1.7.5 Sample handing | Preservation Checks | Requirement | V1M4 1.7.5 b) | The laboratory shall implement procedures for checking sample preservation using readily available techniques, such as pH or chlorine, prior to or during sample preparation or analysis.
1237 | 1.7 Technical Requirement | 1.7.5 Sample handing | Preservation Checks | Exception | V1M4 1.7.5 b) | An exception is allowed for volatile organic compound analyses; chemical preservation may be checked after analysis.

## Module 5
ID | Major Topic | Topic | Subject | Category | Citation | Does the laboratory comply with this section? | Yes | No | N/A
1238 |  |  |  |  | V1M5 1 | MICROBIOLOGICAL TESTING
1239 | 1.1 Introduction |  |  |  |  | Introduction
1240 | 1.1 Introduction |  |  |  | V1M5 1.1 | This Standard applies to laboratories undertaking microbiological analysis of environmental samples. Microbiological testing refers to and includes the detection, isolation, enumeration, or identification of microorganisms and/or their metabolites, or determination of the presence or absence of growth in materials and media. The evaluation of laboratories for this discipline is in conjunction with a quality system as specified in general requirements module. Adherence to quality systems requirements will ensure that all quality control procedures specified in this module are being followed.
1241 | 1.2 Scope |  |  |  | V1M5 1.2 | Scope
1242 | 1.2 Scope |  |  |  | V1M5 1.2 | The essential quality control procedures applicable to microbiological analysis are included in this module. Additional quality control requirements that are either specified by method, regulation or project shall be met by laboratories.
1243 | 1.3 Definitions |  |  |  | V1M5 1.3 | Terms and Definitions
1244 | 1.3 Definitions |  |  |  | V1M5 1.3 | The relevant definitions from TNI, Volume 1, Module 2, Section 3.0 apply. Definitions related to this document, which are used differently or do not exist in the above references are defined below.
1245 | 1.3 Definitions |  |  |  | V1M5 1.3.1 | Additional Terms and Definitions
1246 | 1.3 Definitions |  |  |  | V1M5 1.3.1 | Reserved
1247 | 1.3 Definitions |  |  |  | V1M5 1.3.2 | Exclusions and Exceptions
1248 | 1.3 Definitions |  |  |  | V1M5 1.3.2 | Reserved
1249 | 1.4 Method Selection |  |  |  | V1M5 1.4 | Method Selection
1250 | 1.4 Method Selection | Reference Method | Regulatory | Definition | V1M5 1.4 | A reference method is a method issued by an organization generally recognized as competent to do so. (When ISO refers to a standard method, that term is equivalent to reference method).
1251 | 1.4 Method Selection | Reference Method | Regulatory | Definition | V1M5 1.4 | When a laboratory is required to analyze a parameter by a specified method due to a regulatory requirement, the parameter/method combination is recognized as a reference method.
1252 | 1.4 Method Selection | Reference Method | Regulatory | Definition | V1M5 1.4 | If there is not a regulatory requirement for the parameter/method combination, the parameter/method combination is recognized as a reference method if it can be analyzed by another similar reference method of the same matrix and technology.
1253 | 1.4 Method Selection | Non Reference Method | Customer Communication | Requirement | V1M5 1.4 | When it is necessary to use methods not covered by reference methods, these shall be subject to agreement with the client and shall include a clear specification of the client's requirements and the purpose of the environmental test.
1254 | 1.4 Method Selection | Non Reference Method | Validation | Requirement | V1M5 1.4 | The method developed shall have been validated appropriately before use.
1255 | 1.5 Method Validation |  |  |  | V1M5 1.5 | Method Validation
1256 | 1.5 Method Validation | General |  | Definition | V1M5 1.5 | Validation is the confirmation by examination and the objective evidence that the particular requirements for a specific intended use are fulfilled.
1257 | 1.5 Method Validation | General | Scope | Requirement | V1M5 1.5 | The laboratory shall validate non-reference methods, laboratory-designed/developed methods, reference methods used outside their published scope, and amplifications and modifications of reference methods to confirm that the methods are fit for the intended use.
1258 | 1.5 Method Validation | General | Extent | Requirement | V1M5 1.5 | The validation shall be as extensive as is necessary to meet the needs of the given application or field of application.
1259 | 1.5 Method Validation | General | Documentation | Requirement | V1M5 1.5 | The laboratory shall record the results obtained, the procedure used for the validation, and a statement as to whether the method is fit for the intended use. The minimum requirements for method validation are given in Sections 1.5.1, 1.5.2 and 1.5.3.
1260 | 1.5 Method Validation | General | Record Retention | Requirement | V1M5 1.5 | The laboratory shall maintain documentation of the validation procedure for as long as the method is in use and for at least five (5) years past the date of last use.
1261 | 1.5 Method Validation | General | PT Program | Requirement | V1M5 1.5 | Laboratories shall participate in a proficiency test program when available. The results of these analyses shall be used to evaluate the ability of the laboratory to produce acceptable data.
1262 | 1.5 Method Validation | General | Assessment | Requirement | V1M5 1.5 | The following assessment shall be performed. If no reference method exists, or if the data quality objectives are different from the reference method, then the laboratory shall demonstrate that the method meets the quality objectives for the intended use.
1263 | 1.5 Method Validation | 1.5.1 Accuracy | Procedure | Option | V1M5 1.5.1 | Accuracy – Use at least one (1) known pure reference culture at the anticipated environmental conditions, and compare the method results to that of a reference method.
1264 | 1.5 Method Validation | 1.5.2 Precision | Procedure | Option | V1M5 1.5.2 | Precision – Perform at least ten (10) replicate analyses with both the proposed and reference method, using the target microorganisms of choice. The results shall show that the methods are not statistically different.
1265 | 1.5 Method Validation | 1.5.3 Selectivity | Procedure | Option | V1M5 1.5.3 | Selectivity (sensitivity) – Verify all responses in at least ten (10) samples using mixed cultures that include the target organism(s), and at varying concentrations (microbial identification testing or equivalent processes may be used). Calculate the number of false positive and false negative results.
1266 | 1.6 Demonstration of Capability |  |  |  | V1M5 1.6 | Demonstration of Capability (DOC)
1267 | 1.6 Demonstration of Capability | 1.6.1 General |  |  | V1M5 1.6.1 | General
1268 | 1.6 Demonstration of Capability | 1.6.1 General | Initial DOC | Requirement | V1M5 1.6.1 | Prior to acceptance and institution of any method for data reporting, satisfactory initial DOC is required (see Section 1.6.2).
1269 | 1.6 Demonstration of Capability | 1.6.1 General | Ongoing DOC | Requirement | V1M5 1.6.1 | Thereafter, ongoing DOC (Section 1.6.3), as per the quality control requirements in Section 1.7.3, is required.
1270 | 1.6 Demonstration of Capability | 1.6.1 General | DOC | Exception | V1M5 1.6.1 | In cases where a laboratory analyzes samples using a method that has been in use by the laboratory for at least one year prior to applying for accreditation, and there have been no significant changes in instrument type, personnel or method, the ongoing DOC shall be acceptable as an initial DOC.
1271 | 1.6 Demonstration of Capability | 1.6.1 General | Initial DOC Exception | Requirement | V1M5 1.6.1 | The laboratory shall have records on file to demonstrate that an initial DOC is not required.
1272 | 1.6 Demonstration of Capability | 1.6.1 General | Documentation | Record | V1M5 1.6.1 | For the initial DOC, appropriate records as discussed in Section 1.6.2 shall be completed (1.6.2).
1273 | 1.6 Demonstration of Capability | 1.6.1 General | Initial DOC Frequency | Requirement | V1M5 1.6.1 | An initial DOC shall be completed each time there is a change in instrument type, personnel, or method.(1.6.2)
1274 | 1.6 Demonstration of Capability | 1.6.1 General | Documentation | Requirement | V1M5 1.6.1 | All demonstrations shall be documented.
1275 | 1.6 Demonstration of Capability | 1.6.1 General | Data | Record | V1M5 1.6.1 | All data applicable to the demonstration shall be retained and readily available at the laboratory.
1276 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC |  |  | V1M5 1.6.2 | Initial DOC
1277 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | IDOC | Requirement | V1M5 1.6.2 | An initial DOC shall be made prior to using any method,
1278 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | IDOC Frequency | Requirement | V1M5 1.6.2 | An initial DOC shall be at any time there is a change in instrument type, personnel or method
1279 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | IDOC Frequency | Requirement | V1M5 1.6.2 | An initial DOC shall be made any time that a method has not been performed by the laboratory or analyst in a twelve (12) month period.
1280 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Documentation | Record | V1M5 1.6.2.1 | The laboratory shall document each initial DOC in a manner such that the following information is readily available for each affected employee:
1281 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | DOC Record | Contents | V1M5 1.6.2.1 a) | analyst(s) involved in preparation and/or analysis;
1282 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | DOC Record | Contents | V1M5 1.6.2.1 b) | matrix;
1283 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | DOC Record | Contents | V1M5 1.6.2.1 c) | organism(s);
1284 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | DOC Record | Contents | V1M5 1.6.2.1 d) | identification of method(s) performed;
1285 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | DOC Record | Contents | V1M5 1.6.2.1 e) | identification of laboratory-specific SOP used for analysis, including revision number;
1286 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | DOC Record | Contents | V1M5 1.6.2.1 f) | date(s) of analysis;
1287 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | DOC Record | Contents | V1M5 1.6.2.1 g) | summary of analyses, including information outlined in Section 1.6.2.2.c.
1288 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Requirement | V1M5 1.6.2.2 | If the method or regulation does not specify an initial DOC, the following procedure is acceptable. It is the responsibility of the laboratory to document that other approaches to initial DOC are adequate.
1289 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 a) | The target organism(s) shall be diluted in a volume of clean quality system matrix (a sample in which no target organisms or interferences are present at concentrations that will impact the results of a specific method).
1290 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 a) | This matrix shall be sterile phosphate or sterile peptone solution unless specified by the manufacturer.
1291 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 a) | Prepare at least four (4) aliquots at the concentration specified, or if unspecified, to the countable range for plate methods or working range for most probable number (MPN) type methods.
1292 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 b) | At least four (4) aliquots shall be prepared and analyzed according to the method either concurrently or over a period of days.
1293 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 c) | Using all of the results, convert these results to logarithmic values, then calculate the mean recovery and standard deviation of the log converted results in the appropriate reporting units for each organism of interest. When it is not possible to determine mean and standard deviations, such as for presence/absence, the laboratory shall assess performance against established and documented criteria.
1294 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 d) | For qualitative tests, acceptable performance in a blind study, either internally or externally generated, may be used to meet this Standard, provided that the study consists of a minimum of a blank, a negative culture, and a positive culture for each target organism or metabolite (e.g. b-glucuronidase in E. coli.).
1295 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 e) | Compare the information from c) above to the corresponding acceptance criteria for precision and accuracy in the method (if applicable or in laboratory-generated acceptance criteria (if there are not established mandatory criteria). If all parameters meet the acceptance criteria, the analysis of actual samples may begin. If any one of the parameters does not meet the acceptance criteria, the performance is unacceptable for that parameter.
1296 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 f) | When one or more of the tested parameters fail at least one of the acceptance criteria, the analyst shall proceed according to i) or ii) below.
1297 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 f) i. | Locate and correct the source of the problem and repeat the initial DOC for all parameters of interest beginning with b) above.
1298 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 f) ii. | Repeat the initial DOC for all parameters that failed to meet criteria.
1299 | 1.6 Demonstration of Capability | 1.6.2 Initial DOC | Procedure | Option | V1M5 1.6.2.2 g) | Repeated failure, however, confirms a general problem with the measurement system. If this occurs, locate and correct the source of the problem and repeat the test for all compounds of interest beginning with b).
1300 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC |  |  | V1M5 1.6.3 | Ongoing DOC
1301 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC | Documentation | Procedure | V1M5 1.6.3.1 | The laboratory shall have a documented procedure describing ongoing DOC.
1302 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC | Analyst | Requirement | V1M5 1.6.3.1 | The analyst(s) shall demonstrate ongoing capability by meeting the quality control requirements of the method, laboratory SOP, client specifications, and/or this Standard.
1303 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC | Alternate Procedures | Requirement | V1M5 1.6.3.1 | It is the responsibility of the laboratory to document that other approaches to ongoing DOC are adequate.
1304 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC | Procedure |  | V1M5 1.6.3.2 | This ongoing demonstration may include one of the following or by performing another initial DOC.
1305 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC | Procedure | Option | V1M5 1.6.3.2 a) | Analysis of one sample or clean matrix that is fortified with a known quantity of the target organism, with results meeting the laboratory acceptance criteria for accuracy and, where applicable to the testing technique, also meeting the observational details expected for the presumptive, confirmed and completed phases defined in the method.
1306 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC | Procedure | Option | V1M5 1.6.3.2 b) | Analysis of one sample in duplicate for each target organism and test, with results meeting the laboratory acceptance criterion for precision.
1307 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC | Procedure | Option | V1M5 1.6.3.2 c) | Acceptable results for one-single-blind proficiency test sample for target organisms in each field of accreditation.
1308 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC | Procedure | Option | V1M5 1.6.3.2 d) | Performance of an alternate adequate procedure for the field of accreditation, the procedure and acceptance criteria being documented in the laboratory’s quality system.
1309 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC | Procedure | Option | V1M5 1.6.3.2 e) | A documented process of analyst review using QC samples. QC samples can be reviewed to identify patterns for individuals or groups of analysts and determine if corrective action or retraining is necessary; or
1310 | 1.6 Demonstration of Capability | 1.6.3 Ongoing DOC | Procedure | Option | V1M5 1.6.3.2 f) | f a) through e) are not technically feasible, then analysis of real-world samples with results within a predefined acceptance criteria (as defined by the laboratory or method) shall be performed.
1311 | 1.7 Technical Requirements |  |  |  | V1M5 1.7 | Technical Requirements
1312 | 1.7 Technical Requirements | 1.7.1 Calibration |  |  | V1M5 1.7.1 | Calibration
1313 | 1.7 Technical Requirements | 1.7.1 Calibration | Calibration | Procedure | V1M5 1.7.1 a) | The laboratory shall have documented procedures for calibration, verification, and quality control of support equipment including conductivity meters, oxygen meters, pH meters, hygrometers, and other similar measurement instruments.
1314 | 1.7 Technical Requirements | 1.7.1 Calibration | Procedure Origin | Requirement | V1M5 1.7.1 a) | These procedures shall refer to applicable reference methods.
1315 | 1.7 Technical Requirements | 1.7.1 Calibration | Continuous Monitors |  | V1M5 1.7.1 b) | For instruments that are continuous monitors, such as in-line specific conductance meters:
1316 | 1.7 Technical Requirements | 1.7.1 Calibration | Continuous Monitors | Requirement | V1M5 1.7.1 b) i. | The laboratory shall document acceptable calibration verification at least once a month.
1317 | 1.7 Technical Requirements | 1.7.1 Calibration | Continuous Monitors | Requirement | V1M5 1.7.1 b) ii. | An initial calibration shall be performed if a continuing calibration is unacceptable,
1318 | 1.7 Technical Requirements | 1.7.1 Calibration | Continuous Monitors | Requirement | V1M5 1.7.1 b) ii. | An initial calibration shall be performed when the instrument is being returned to service after having been taken off line.
1319 | 1.7 Technical Requirements | 1.7.2 Continuing Calibration |  |  | V1M5 1.7.2 | Continuing Calibration
1320 | 1.7 Technical Requirements | 1.7.2 Continuing Calibration |  |  | V1M5 1.7.2 | Reserved for specific procedures.
1321 | 1.7 Technical Requirements | 1.7.3 Quality Control |  |  | V1M5 1.7.3 | Quality Control
1322 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks and Method Blanks |  |  | V1M5 1.7.3.1 | Sterility Checks and Method Blanks
1323 | 1.7 Technical Requirements | 1.7.3.1 Method Blanks |  |  | V1M5 1.7.3.1 a) | Method Blanks
1324 | 1.7 Technical Requirements | 1.7.3.1 Method Blanks | Performance Criteria | Requirement | V1M5 1.7.3.1 a) | The laboratory shall demonstrate that the filtration equipment and filters, sample containers, media and reagents have not been contaminated through improper handling or preparation, inadequate sterilization, or environmental exposure.
1325 | 1.7 Technical Requirements | 1.7.3.1 Method Blanks | Filtration | Requirement | V1M5 1.7.3.1 a) i. | For filtration technique, the laboratory shall conduct method blanks per the analytical method.
1326 | 1.7 Technical Requirements | 1.7.3.1 Method Blanks | Filtration | Requirement | V1M5 1.7.3.1 a) i. | At a minimum, the filtration series shall include a beginning and ending blank.
1327 | 1.7 Technical Requirements | 1.7.3.1 Method Blanks | Filtration | Requirement | V1M5 1.7.3.1 a) i. | The filtration series may include single or multiple filtration units, which have been sterilized prior to beginning the series.
1328 | 1.7 Technical Requirements | 1.7.3.1 Method Blanks | Filtration Series | Requirement | V1M5 1.7.3.1 a) ii. | The filtration series is considered ended when more than thirty (30) minutes elapses between successive filtrations.
1329 | 1.7 Technical Requirements | 1.7.3.1 Method Blanks | Filtration Procedure | Requirement | V1M5 1.7.3.1 a) ii. | During a filtration series, filter funnels shall be rinsed with three (3) 20-30 ml portions of sterile rinse water after each sample filtration.
1330 | 1.7 Technical Requirements | 1.7.3.1 Method Blanks | Filtration Procedure | Requirement | V1M5 1.7.3.1 a) ii. | In addition, laboratories shall insert a method blank after every ten (10) samples or sanitize filtration units by UV light after each sample filtration.
1331 | 1.7 Technical Requirements | 1.7.3.1 Method Blanks | Pour Place | Requirement | V1M5 1.7.3.1 a) iii. | For pour plate technique, method blanks of the medium shall be made by pouring, at a minimum, one uninoculated plate for each lot of pre-prepared, ready-to-use media
1332 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks |  |  | V1M5 1.7.3.1 b) | Sterility Checks
1333 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Prepared Media | Requirement | V1M5 1.7.3.1 b) i. | A sterility check shall be analyzed for each lot of pre-prepared, ready-to-use medium (including chromofluorogenic reagent)
1334 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Laboratory Prepared Media | Requirement | V1M5 1.7.3.1 b) i. | A sterility check shall be analyzed for each batch of medium prepared in the laboratory.
1335 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Frequency | Requirement | V1M5 1.7.3.1 b) i. | A sterility check shall be done prior to first use of the medium.
1336 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Presterilized Funnels | Requirement | V1M5 1.7.3.1 b) ii. | For pre-sterilized single use funnels, a sterility check shall be performed on one funnel per lot
1337 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Laboratory Sterilized Funnels | Requirement | V1M5 1.7.3.1 b) ii. | For laboratory-sterilized funnels, a sterility check shall be performed on one funnel per sterilization batch.
1338 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Sample Containers | Requirement | V1M5 1.7.3.1 b) iii. | Sterility checks on sample containers shall be performed on at least one (1) container for each lot of purchased, pre-sterilized containers.
1339 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Laboratory Sterilized Containers | Requirement | V1M5 1.7.3.1 b) iii. | For containers prepared and sterilized in the laboratory, a sterility check shall be performed on one (1) container per sterilized batch with nonselective growth media.
1340 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Contract Laboratory | Option | V1M5 1.7.3.1 b) iii. | These sterility checks may be performed by a contracted laboratory if the laboratory does not have the requisite equipment to perform them.
1341 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Record Retention | Requirement | V1M5 1.7.3.1 b) iii. | All correspondence and results from a contracted laboratory shall be retained for a period of five (5) years after the completion of the test(s).
1342 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Dilution Water | Requirement | V1M5 1.7.3.1 b) iv. | A sterility check shall be performed on each batch of dilution water prepared in the laboratory with non-selective growth media.
1343 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Purchased Dilution Water | Requirement | V1M5 1.7.3.1 b) iv. | A sterility check shall be performed on each lot of pre-prepared, ready-to-use dilution water with non-selective growth media.
1344 | 1.7 Technical Requirements | 1.7.3.1 Sterility Checks | Membrane Filter | Requirement | V1M5 1.7.3.1 b) v. | At least one (1) filter from each new lot of membrane filters shall be checked for sterility with nonselective growth media.
1345 | 1.7 Technical Requirements | 1.7.3.2 Test Variability |  |  | V1M5 1.7.3.2 | Test Variability/Reproducibility
1346 | 1.7 Technical Requirements | 1.7.3.2 Test Variability | Duplicate Counts | Requirement | V1M5 1.7.3.2 | For methods that specify colony counts such as membrane filter or plated media, duplicate counts shall be performed monthly on one positive sample, for each month that the test is performed. If the laboratory has two or more analysts, each analyst shall count typical colonies on the same plate. Counts shall be within 10% difference to be acceptable. In a laboratory with only one microbiology analyst, the same plate shall be counted twice by the analyst, with no more than 5% difference between the counts.
1347 | 1.7 Technical Requirements | 1.7.3.3 Sample Specific Controls |  |  | V1M5 1.7.3.3 | Sample Specific Controls (where applicable)
1348 | 1.7 Technical Requirements | 1.7.3.3 Sample Specific Controls | Matrix Spike | Requirement | V1M5 1.7.3.3 a) | Matrix spikes shall be performed per method requirements.
1349 | 1.7 Technical Requirements | 1.7.3.3 Sample Specific Controls | Sample Matrix Duplicates | Requirement | V1M5 1.7.3.3 b) | Sample matrix duplicates shall be performed per method requirements.
1350 | 1.7 Technical Requirements | 1.7.3.4 Data Reduction |  |  | V1M5 1.7.3.4 | Data Reduction
1351 | 1.7 Technical Requirements | 1.7.3.4 Data Reduction | Calculations and Statistics | Requirement | V1M5 1.7.3.4 | The calculations, data reduction and statistical interpretations specified by each method shall be identified and followed.
1352 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality |  |  | V1M5 1.7.3.5 | Quality of Standards, Reagents and Media
1353 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Appropriateness | Requirement | V1M5 1.7.3.5 | The laboratory shall ensure that the quality of the reagents and media used is appropriate for the test concerned.
1354 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Media Preparation | Option | V1M5 1.7.3.5 a) | Media – Culture media may be prepared from commercial dehydrated powders or may be purchased ready-to-use.
1355 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Laboratory Prepared Media |  | V1M5 1.7.3.5 a) i. | Laboratory-prepared media
1356 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Laboratory Prepared Media | Requirement | V1M5 1.7.3.5 a) i. 1 | Media prepared by the laboratory from basic ingredients shall be tested for performance (e.g., for selectivity, sensitivity, sterility, growth promotion, and growth inhibition) prior to first use.
1357 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Laboratory Prepared Media | Requirement | V1M5 1.7.3.5 a) i. 2 | Media shall be used within the holding time limits specified in the accredited method.
1358 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Laboratory Prepared Media | Requirement | V1M5 1.7.3.5 a) i. 3 | Detailed testing criteria information shall be defined in the laboratory’s methods, SOPs, or similar documentation.
1359 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Ready to Use Media |  | V1M5 1.7.3.5 a) ii. | Ready-to-use media
1360 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Ready to Use Media | Requirement | V1M5 1.7.3.5 a) ii. 1 | Ready-to-use media shall be used within the manufacturer’s expiration date. If the manufacturer’s expiration date is greater than those noted in Section 1.7.3.5 a) i) 2. above, the laboratory shall request, and have available documentation from the manufacturer demonstrating media quality for the extended time period.
1361 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Ready to Use Media | Requirement | V1M5 1.7.3.5 a) ii. 2 | Any ready-to-use media used past the expiration date shall be verified for usability by running quality control checks comparing the media with freshly prepared media or by testing recovery with known densities of culture controls.
1362 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Ready to Use Media | Requirement | V1M5 1.7.3.5 b) | Reagents and commercial dehydrated powders shall be used within the shelf life of the product, and shall be documented as per TNI Volume 1, Module 2 Quality Systems General Requirements.
1363 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Reagent Water |  | V1M5 1.7.3.5 c) | Reagent Water
1364 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Reagent Water | Requirement | V1M5 1.7.3.5 c) i. | The quality of the reagent water used in the laboratory, such as distilled water, de-ionized water or reverse-osmosis produced water shall be monitored for bactericidal and inhibitory substances and shall be used in the preparation of media, solutions and buffers.
1365 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Reagent Water | Requirement | V1M5 1.7.3.5 c) ii. | The quality of the water shall be monitored for chlorine residual, specific conductance, total organic carbon, ammonia/organic nitrogen and heterotrophic bacteria plate count monthly (when in use), when maintenance is performed on the water treatment system, or at startup after a period of disuse longer than one month.
1366 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Reagent Water | Requirement | V1M5 1.7.3.5 c) iii. | Analysis for metals and the Bacteriological Water Quality Test (to determine presence of toxic agents or growth promoting substances) shall be performed annually. (An exception to performing the Bacteriological Water Quality Test shall be given to laboratories that can supply documentation to show that their water source meets the criteria, as specified by the method, for Type I or Type II reagent water.)
1367 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Reagent Water | Record | V1M5 1.7.3.5 c) iv. | Results of the above analyses (metals and bacteriological water quality) shall meet the specifications of the required method and records of analyses shall be maintained for five (5) years.
1368 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Reagent Water | Requirement | V1M5 1.7.3.5 c) v. | Reagent water purchased from an outside source and used for the preparations of media, solutions and buffers shall meet the criteria specified in items ii) and iii) above. The laboratory shall have documented records of this information. Purchased reagent water that has been in use for longer than the testing intervals specified in items i) through i or in the accredited method shall either be re-tested or discarded.
1369 | 1.7 Technical Requirements | 1.7.3.5 Standard, Reagent & Media Quality | Reagent Water | Requirement | V1M5 1.7.3.5 d) | Documentation for media prepared in the laboratory shall include date of preparation, preparer’s initials, type , manufacturer, lot number, final pH, expiration date, and the amount of reagents used. Documentation for media purchased pre-prepared, ready-to-use (including reagent water purchased from outside sources) shall include manufacturer, lot number, type of media received, date of receipt, expiration date of the media, and pH of the media.
1370 | 1.7 Technical Requirements | 1.7.3.6 Selectivity |  |  | V1M5 1.7.3.6 | Selectivity
1371 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Media checks | Requirement | V1M5 1.7.3.6 a) | All growth and recovery media shall be checked to assure that the target organism(s) respond in an acceptable and predictable manner.
1372 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Organism Verification | Requirement | V1M5 1.7.3.6 b) | To ensure that analysis results are accurate, target organism identity shall be verified as specified in the method (e.g., by use of the completed test, or by use of secondary verification tests such as a catalase test or by the use of a completed test such as brilliant green (BG) or E. coli (EC) broth.
1373 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Reference Cultures | Requirement | V1M5 1.7.3.6 c) | In order to ensure identity and traceability, reference cultures used for positive and negative controls shall be obtained from a recognized national collection, organization, or manufacturer recognized by the accreditation body.
1374 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Reference Cultures | Option | V1M5 1.7.3.6 c) | Microorganisms may be single use preparations or cultures maintained for their intended use by documented procedures that demonstrate the continued purity and viability of the organism.
1375 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Reference Cultures | Requirement | V1M5 1.7.3.6 c) i. | Reference cultures may be revived (if freeze-dried) or transferred from slants and sub-cultured once to provide reference stocks. The reference stocks shall be preserved by a technique that maintains the characteristics of the strains. Reference stocks shall be used to prepare working stocks for routine work. If reference stocks have been thawed, they shall not be refrozen and re-used.
1376 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Working Stocks | Requirement | V1M5 1.7.3.6 c) ii. | Working stocks shall not be sequentially cultured more than five (5) times and shall not be sub-cultured to replace reference stocks.
1377 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Culture Controls |  | V1M5 1.7.3.6 d) | Culture Controls
1378 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Negative Culture Control |  | V1M5 1.7.3.6 d) i. | Negative Culture Controls
1379 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Negative Culture Control |  | V1M5 1.7.3.6 d) i. 1 | Negative culture controls demonstrate that the medium does not support the growth of non-target organisms or does not exhibit the typical positive reaction of the target organism(s).
1380 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Negative Culture Control | Requirement | V1M5 1.7.3.6 d) i. 2 | Each pre-prepared, ready-to-use lot of selective medium (including chromofluorogenic reagent) and each batch of selective medium prepared in the laboratory shall be analyzed with one or more known negative culture controls (i.e. non-target organisms), as appropriate to the method. This shall be done prior to first use of the medium.
1381 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Positive Culture Control |  | V1M5 1.7.3.6 d) ii. | Positive Culture Controls
1382 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Positive Culture Control |  | V1M5 1.7.3.6 d) ii. 1 | Positive culture controls demonstrate that the medium can support the growth of the target organism(s), and that the medium produces the specified or expected reaction to the target organism(s).
1383 | 1.7 Technical Requirements | 1.7.3.6 Selectivity | Positive Culture Control | Requirement | V1M5 1.7.3.6 d) ii. 2 | Each pre-prepared, ready-to-use lot of medium (including chromofluorogenic reagent) and each batch of medium prepared in the laboratory shall be tested with at least one pure culture of a known positive reaction. This shall be done prior to first use of the medium.
1384 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions |  |  | V1M5 1.7.3.7 | Constant and Consistent Test Conditions
1385 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Laboratory Facilities |  | V1M5 1.7.3.7 a) | Laboratory Facilities
1386 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Laboratory Facilities | Requirement | V1M5 1.7.3.7 a) | Floors and work surfaces shall be non-absorbent and easy to clean and disinfect. Work surfaces shall be adequately sealed.
1387 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Laboratory Facilities | Requirement | V1M5 1.7.3.7 a) | Laboratories shall provide sufficient storage space, and shall be clean and free from dust accumulation.
1388 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Laboratory Facilities | Requirement | V1M5 1.7.3.7 a) | Plants, food, and drink shall be prohibited from the laboratory work area.
1389 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Laboratory Equipment |  | V1M5 1.7.3.7 b) | Laboratory Equipment
1390 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Temperature Measuring Devices |  | V1M5 1.7.3.7 b) i. | Temperature Measuring Devices
1391 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Temperature Measuring Devices | Requirement | V1M5 1.7.3.7 b) i. | Temperature measuring devices such as liquid-in-glass thermometers, thermocouples, and platinum resistance thermometers used in incubators, autoclaves and other equipment shall be the appropriate quality to meet specification(s) in the method.
1392 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Temperature Measuring Devices | Requirement | V1M5 1.7.3.7 b) i. | The graduation of the temperature measuring devices shall be appropriate for the required accuracy of measurement and they shall be verified to national or international standards for temperature.
1393 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Temperature Measuring Devices | Requirement | V1M5 1.7.3.7 b) i. | Verification shall be done at least annually (see TNI Volume 1, Module 2, Section 5.5.13.1).
1394 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Autoclaves |  | V1M5 1.7.3.7 b) ii. | Autoclaves
1395 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Autoclaves | Requirement | V1M5 1.7.3.7 b) ii. | The performance of each autoclave shall be initially evaluated by establishing its functional properties and performance, for example heat distribution characteristics with respect to typical uses.
1396 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Autoclaves | Requirement | V1M5 1.7.3.7 b) ii. | Autoclaves shall meet specified temperature tolerances. Pressure cookers shall not be used for sterilization of growth media.
1397 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Autoclaves | Requirement | V1M5 1.7.3.7 b) ii. | Pressure cookers shall not be used for sterilization of growth media.
1398 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Sterilization Temperature | Requirement | V1M5 1.7.3.7 b) ii. | Demonstration of sterilization temperature shall be provided by use of a continuous temperature recording device or by use of a maximum registering thermometer with every cycle.
1399 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Sterilization Temperature | Requirement | V1M5 1.7.3.7 b) ii. | At least once during each month that the autoclave is used, appropriate biological indicators shall be used to determine effective sterilization. The selected biological indicator shall be effective at the sterilization temperature and time needed to sterilize lactose-based media.
1400 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Sterilization Temperature | Requirement | V1M5 1.7.3.7 b) ii. | Temperature sensitive tape shall be used with the contents of each autoclave run to indicate that the autoclave contents have been processed.
1401 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Autoclaves | Record | V1M5 1.7.3.7 b) ii. | Records of autoclave operations shall be maintained for every cycle. Records shall include: date, contents, maximum temperature reached, pressure, time in sterilization mode, total run time (may be recorded as time in and time out) and analyst’s initials.
1402 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Autoclave Maintenance | Requirement | V1M5 1.7.3.7 b) ii. | Autoclave maintenance, either internally or by service contract, shall be performed annually, and shall include a pressure check and verification of temperature device. Records of the maintenance shall be maintained in equipment logs.
1403 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Autoclave Maintenance | Option | V1M5 1.7.3.7 b) ii. | NOTE: When it has been determined that the autoclave has no leaks, pressure checks can be documented using the formula PV = nRT.
1404 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Autoclave Mechanical Timing Device | Requirement | V1M5 1.7.3.7 b) ii. | The autoclave mechanical timing device shall be checked quarterly against a stopwatch and the actual time elapsed documented.
1405 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Volumetric Equipment |  | V1M5 1.7.3.7 b) iii. | Volumetric Equipment
1406 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Volumetric Equipment |  | V1M5 1.7.3.7 b) iii. | Volumetric equipment shall be verified as follows:
1407 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Volumetric Equipment Verification | Requirement | V1M5 1.7.3.7 b) iii. 1 | equipment with movable parts such as automatic dispensers, dispensers/diluters, and mechanical hand pipettes shall be verified for accuracy quarterly.
1408 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Volumetric Equipment Verification | Requirement | V1M5 1.7.3.7 b) iii. 2 | equipment such as filter funnels, bottles, non-Class A glassware, and other containers with volumetric markings (including sample analysis vessels) shall be verified once per lot prior to first use. This verification may be volumetric or gravimetric.
1409 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Volumetric Equipment Verification | Requirement | V1M5 1.7.3.7 b) iii. 3 | the volume of the disposable volumetric equipment such as sample bottles, and disposable pipettes shall be checked once per lot.
1410 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | UV Instruments |  | V1M5 1.7.3.7 b) iv. | UV Instruments
1411 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | UV Instruments | Requirement | V1M5 1.7.3.7 b) iv. | UV instruments, used for sanitization, shall be tested quarterly for effectiveness with an appropriate UV light meter, by plate count agar spread plates or other methods providing equivalent results such as uvcide strips.
1412 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | UV Instruments | Requirement | V1M5 1.7.3.7 b) iv. | Replace bulbs if output is less than 70% of original for light tests or if count reduction is less than 99% for a plate containing 200 to 300 organisms.
1413 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Incubators, Water Baths, Ovens |  | V1M5 1.7.3.7 b) v. | Incubators, Water Baths, Ovens
1414 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Incubators, Water Baths, Ovens | Requirement | V1M5 1.7.3.7 b) v. 1 | The uniformity of temperature distribution in incubators and water baths shall be established.
1415 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Incubators, Water Baths, Ovens | Record | V1M5 1.7.3.7 b) v. 2 | Records shall be maintained for each cycle that include date, cycle time, temperature, contents and analyst’s initials.
1416 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Labware |  | V1M5 1.7.3.7 b) vi. | Labware (Glassware and Plasticware)
1417 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Labware | Procedure | V1M5 1.7.3.7 b) vi. 1 | The laboratory shall have a documented procedure for washing labware, if applicable. Detergents designed for laboratory use shall be used.
1418 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Labware | Requirement | V1M5 1.7.3.7 b) vi. 2 | Glassware shall be made of borosilicate or other non-corrosive material, free of chips and cracks, and shall have readable measurement marks.
1419 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Labware Inhibitory Residue | Requirement | V1M5 1.7.3.7 b) vi. 3 | Labware that is washed and reused shall be tested for possible presence of residues that may inhibit or promote growth of microorganisms by performing the Inhibitory Residue Test annually, and each time the lab changes the lot of detergent or washing procedures.
1420 | 1.7 Technical Requirements | 1.7.3.7 Test Conditions | Labware Residue | Requirement | V1M5 1.7.3.7 b) vi. 4 | Washed labware shall be tested at least once daily, each day of washing, for possible acid or alkaline residue by testing at least one piece of labware with a suitable pH indicator such as bromothymol blue. Records of tests shall be maintained.
1421 | 1.7 Technical Requirements | 1.7.4 Data Acceptance / Rejection |  |  | V1M5 1.7.4 | Data Acceptance/Rejection Criteria
1422 | 1.7 Technical Requirements | 1.7.4 Data Acceptance / Rejection | Method Procedures | Requirement | V1M5 1.7.4 | Method criteria and evaluation methods shall be used.
1423 | 1.7 Technical Requirements | 1.7.5 Sample Handling |  |  | V1M5 1.7.5 | Sample Handling
1424 | 1.7 Technical Requirements | 1.7.5 Sample Handling | Thermal Preservation | Requirement | V1M5 1.7.5 a) | Samples that require thermal preservation shall be considered acceptable if the arrival temperature of a representative sample container meets the method or mandated temperature requirement.
1425 | 1.7 Technical Requirements | 1.7.5 Sample Handling | Thermal Preservation | Option | V1M5 1.7.5 a) i. | Samples that are delivered to the laboratory on the same day they are collected may not meet the requirements of Section 1.7.5.a). In these cases, the samples shall be considered acceptable if the samples were received on ice.
1426 | 1.7 Technical Requirements | 1.7.5 Sample Handling | Thermal Preservation | Option | V1M5 1.7.5 a) ii. | If sample analysis is begun within fifteen (15) minutes of collection, thermal preservation is not required.
1427 | 1.7 Technical Requirements | 1.7.5 Sample Handling | Thermal Preservation | Option | V1M5 1.7.5 a) iii. | Thermal preservation is not required in the field if the laboratory receives the sample and either begins the analysis or refrigerates the sample within fifteen (15) minutes of collection.
1428 | 1.7 Technical Requirements | 1.7.5 Sample Handling | Chlorine Checks |  | V1M5 1.7.5 b) | Microbiological samples from known chlorinated sources (such as wastewater effluent), unknown sources where chlorine usage is suspected (such a new client or a new source) and all potable water sources (including source water) shall be checked for absence of chlorine residual.
1429 | 1.7 Technical Requirements | 1.7.5 Sample Handling | Chlorine Checks |  | V1M5 1.7.5 b) | Laboratories that receive samples from potable water sources (including source water) that have a demonstrated history of acceptable preservation may check a sample from each source at a frequency of once per month if:
1430 | 1.7 Technical Requirements | 1.7.5 Sample Handling | Chlorine Checks | Option | V1M5 1.7.5 b) i. | the laboratory can show that the received sample containers are from their laboratory;
1431 | 1.7 Technical Requirements | 1.7.5 Sample Handling | Chlorine Checks | Option | V1M5 1.7.5 b) ii. | sufficient sodium thiosulfate was in each container before sample collection to neutralize at minimum 5 mg/l of chlorine for drinking water and 15 mg/l of chlorine for wastewater samples;
1432 | 1.7 Technical Requirements | 1.7.5 Sample Handling | Chlorine Checks | Option | V1M5 1.7.5 b) iii. | one container from each batch of laboratory prepared containers or lot of purchased ready-to-use containers is checked to ensure efficacy of the sodium thiosulfate to 5 mg/l chlorine or 15 mg/l chlorine as appropriate and the check is documented;
1433 | 1.7 Technical Requirements | 1.7.5 Sample Handling | Chlorine Checks | Option | 1.7.5 b) iv. | chlorine residual is checked in the field and actual concentration is documented with sample submission.

## Module 6
ID | Major Topic | Topic | Subject | Category | Citation | Does the laboratory comply with this section? | Yes | No | N/A
1434 | 1.0 Radiochemical testing |  |  |  | V1M6 1 | 1.0 RADIOCHEMICAL TESTING
1435 | 1.1 Introduction |  |  |  | V1M6 1.1 | Introduction
1436 | 1.1 Introduction |  |  |  | V1M6 1.1 | This Standard contains detailed quality control requirements for environmental testing activities involving radiochemical measurements. The evaluation of laboratories for this discipline is in conjunction with a quality system as specified in the general requirements module. Adherence to quality systems requirements will ensure that all quality control procedures specified in this module are being followed.
1437 | 1.2 Scope |  |  |  | V1M6 1.2 | Scope
1438 | 1.2 Scope |  |  |  | V1M6 1.2 | These requirements apply to laboratories undertaking the examination of environmental samples by radiochemical analysis. Procedures for radiochemical analysis may involve some form of chemical separation followed by detection of the radioactive emissions of the analyte (or indicative daughters) and tracer isotopes where used. Procedures for the determination of radioactive isotopes by mass spectrometry (e.g., ICP-MS or TIMS) or optical (e.g., KPA) techniques are outside the scope of this document.
1439 | 1.2 Scope |  |  |  | V1M6 1.2 | The essential quality control procedures applicable to radiochemistry measurements are included in this Standard. Additional quality control requirements that are specified by method, regulation or project shall be met by laboratories.
1440 | 1.3 Definitions |  |  |  | V1M6 1.3 | Terms and Definitions
1441 | 1.3 Definitions |  |  |  | V1M6 1.3 | The relevant definitions from TNI, Volume 1, Module 2, Section 3.0 apply. Definitions related to this document, which are used differently or do not exist in the above references are defined below.
1442 | 1.3 Definitions |  |  |  | V1M6 1.3.1 | Additional Terms and Definitions
1443 | 1.3 Definitions |  |  |  | V1M6 1.3.1 | Reserved
1444 | 1.3 Definitions |  |  |  | V1M6 1.3.2 | Exclusions and Exceptions
1445 | 1.3 Definitions |  |  |  | V1M6 1.3.2 | Reserved
1446 | 1.4 Method Selection |  |  |  | V1M6 1.4 | Method Selection
1447 | 1.4 Method Selection | Reference Method | Regulatory | Definition | V1M6 1.4 | A reference method is a method issued by an organization generally recognized as competent to do so. (When ISO refers to a standard method, that term is equivalent to reference method).
1448 | 1.4 Method Selection | Reference Method | Regulatory | Definition | V1M6 1.4 | When a laboratory is required to analyze a parameter by a specific method due to a regulatory requirement, the parameter/method combination is recognized as a reference method.
1449 | 1.4 Method Selection | Reference Method | QC | Requirement | V1M6 1.4 | If there is not a regulatory requirement for the parameter/method combination, the parameter/method combination is recognized as a reference method if it can be analyzed by another similar reference method of the same matrix and technology, and the inclusion of the parameter in the method meets all required calibration requirements of the method and the quality control requirements of the method to which the parameter is being added. If no QC exists in the method, the laboratory shall adhere to the requirements outlined in the similar method.
1450 | 1.4 Method Selection | Non Reference Method | Customer Communication | Requirement | V1M6 1.4 | When it is necessary to use methods not covered by reference methods, these shall be subject to agreement with the client and shall include a clear specification of the client's requirements and the purpose of the environmental test.
1451 | 1.4 Method Selection | Non Reference Method | Validation | Requirement | V1M6 1.4 | The method developed shall have been validated appropriately before use.
1452 | 1.5 Method Validation |  |  |  | V1M6 1.5 | Method Validation
1453 | 1.5 Method Validation | 1.5.1 Validation |  |  | V1M6 1.5.1 | Validation of Methods
1454 | 1.5 Method Validation | 1.5.1 Validation | General | Definition | V1M6 1.5.1 a) | Validation is the confirmation by examination and the objective evidence that the particular requirements for a specific intended use are fulfilled.
1455 | 1.5 Method Validation | 1.5.1 Validation | Reference Method | Requirement | V1M6 1.5.1 b) | For reference methods, the minimum detectable activity (Section 1.5.2.1) applies. Evaluating precision and bias is covered in Section 1.5.3.
1456 | 1.5 Method Validation | 1.5.1 Validation | Non Reference Method Validation | Requirement | V1M6 1.5.1 c) | The laboratory shall validate non-reference methods, laboratory-designed/developed methods, reference methods used outside their published scope, and amplifications and modifications of reference methods to confirm that the methods are fit for the intended use.
1457 | 1.5 Method Validation | 1.5.1 Validation | Non Reference Method Validation | Requirement | V1M6 1.5.1 c) | The validation shall be as extensive as is necessary to meet the needs of the given application or field of application.
1458 | 1.5 Method Validation | 1.5.1 Validation | Documentation | Requirement | V1M6 1.5.1 c) | The laboratory shall record the results obtained, the procedure used for the validation, and a statement as to whether the method is fit for the intended use. The minimum requirements for method validation are given in Sections 1.5.2 – 1.5.5.
1459 | 1.5 Method Validation | 1.5.2 Detectable Activity |  |  | V1M6 1.5.2 | Detectable Activity
1460 | 1.5 Method Validation | 1.5.2 Detectable Activity | General | Procedure | V1M6 1.5.2 | Detectable Activity - All procedures used shall be documented.
1461 | 1.5 Method Validation | 1.5.2 Detectable Activity | Documentation | Requirement | V1M6 1.5.2 | Detectable Activity - Documentation shall include the quality system matrix type.
1462 | 1.5 Method Validation | 1.5.2 Detectable Activity | Data | Record | V1M6 1.5.2 | Detectable Activity - All supporting data shall be retained.
1463 | 1.5 Method Validation | 1.5.2.1 Minimum Detectable Activity |  |  | V1M6 1.5.2.1 | Minimum Detectable Activity (MDA)
1464 | 1.5 Method Validation | 1.5.2.1 Minimum Detectable Activity | Relevancy | Procedure | V1M6 1.5.2.1 | The laboratory shall utilize a method that provides an MDA that is appropriate and relevant for the intended use of the data.
1465 | 1.5 Method Validation | 1.5.2.1 Minimum Detectable Activity | Mandated Method | Requirement | V1M6 1.5.2.1 | MDAs shall be determined by the protocol in the mandated method.
1466 | 1.5 Method Validation | 1.5.2.1 Minimum Detectable Activity | Selection Criteria | Procedure | V1M6 1.5.2.1 | If the protocol for determining the MDA is not specified, the selection of the procedure shall reflect instrument limitations and the intended application of the method.
1467 | 1.5 Method Validation | 1.5.2.1 Minimum Detectable Activity | Target Analytes | Requirement | V1M6 1.5.2.1 a) | The laboratory shall determine the MDA for the method for each target analyte of concern in the quality system sample matrices.
1468 | 1.5 Method Validation | 1.5.2.1 Minimum Detectable Activity | Processing | Requirement | V1M6 1.5.2.1 a) | All sample-processing steps of the analytical method shall be included in the determination of the MDA.
1469 | 1.5 Method Validation | 1.5.2.1 Minimum Detectable Activity | Quality System Matrix | Requirement | V1M6 1.5.2.1 b) | The MDA shall be initially determined for the analytes of interest in each method in a quality system matrix in which there are no target analytes and no interferences at levels that would impact the results.
1470 | 1.5 Method Validation | 1.5.2.1 Minimum Detectable Activity | Frequency | Requirement | V1M6 1.5.2.1 c) | The MDA shall be determined each time there is a change in the method that affects how the test is performed
1471 | 1.5 Method Validation | 1.5.2.1 Minimum Detectable Activity | Frequency | Requirement | V1M6 1.5.2.1.c) | The MDA shall be determined each time a change in instrumentation occurs that affects the analytical detection capability.
1472 | 1.5 Method Validation | 1.5.2.1 Minimum Detectable Activity | MDA Definition | Requirement | V1M6 1.5.2.1 d) | The MDA is an estimate of the smallest true activity (or activity concentration) of analyte in a sample that ensures a 95% probability of detection, given a detection criterion that ensures only a 5% probability of detection in analyte-free samples.
1473 | 1.5 Method Validation | 1.5.2.2 Detection Limits for SDWA |  |  | V1M6 1.5.2.2 | Required Detection Limit for Drinking Water Compliance
1474 | 1.5 Method Validation | 1.5.2.2 Detection Limits for SDWA | Method Selection | Requirement | V1M6 1.5.2.2 | Laboratories that analyze drinking-water samples for Safe Drinking Water Act (SDWA) compliance monitoring shall use methods whose detection limits meet the requirements of 40 CFR 141.
1475 | 1.5 Method Validation | 1.5.2.2 Detection Limits for SDWA | Detection Limit Definition | Requirement | V1M6 1.5.2.2 | The SDWA detection limit is defined in 40 CFR 141.25(c) as equal to the analyte concentration which can be counted with a precision of plus or minus 100% at the 95% confidence level (1.96σ where σ is the standard deviation of the net counting rate of the sample). The SDWA detection limit is equivalent to the concentration at which the relative standard deviation of the measurement due to counting statistics is 1/1.96.
1476 | 1.5 Method Validation | 1.5.3 Precision and Bias |  |  | V1M6 1.5.3 | Evaluation of Precision and Bias
1477 | 1.5 Method Validation | 1.5.3 Precision and Bias | Reference Method | Requirement | V1M6 1.5.3 a) | Reference Methods. The laboratory shall evaluate the precision and bias of a reference method for each analyte of concern for each quality system matrix according to Section 1.6 or alternate documented procedure when the analyte cannot be spiked into the sample matrix and QC samples are not commercially available.
1478 | 1.5 Method Validation | 1.5.3 Precision and Bias | Non Reference Method | Procedure | V1M6 1.5.3 b) | Non- Reference Methods. For laboratory-developed methods or non-reference methods that were not in use by the laboratory before July 2003, the laboratory shall have a documented procedure to evaluate precision and bias.
1479 | 1.5 Method Validation | 1.5.3 Precision and Bias | Evaluation | Procedure | V1M6 1.5.3.b) | The laboratory shall also compare results of the precision and bias measurements with criteria established by the client, given in the reference method, or established by the laboratory.
1480 | 1.5 Method Validation | 1.5.3 Precision and Bias | Evaluation | Requirement | V1M6 1.5.3 c) | The laboratory shall also evaluate precision and bias in the relevant quality system matrices
1481 | 1.5 Method Validation | 1.5.3 Precision and Bias | Processing | Requirement | V1M6 1.5.3 c) | The laboratory shall process the samples through the entire measurement system for each analyte of interest.
1482 | 1.5 Method Validation | 1.5.3 Precision and Bias | Example |  | V1M6 1.5.3 d) | An example of a systematic approach to evaluate precision and bias could be the following:
1483 | 1.5 Method Validation | 1.5.3 Precision and Bias | Example | Option | V1M6 1.5.3 d) | Analyze QC samples in triplicate containing the analytes of concern at or near the MDA, at a level near ten (10) times the MDA, and at a mid-range concentration. Process these samples on different days as three (3) sets of samples through the entire measurement system for each analyte of interest. Each day one QC sample at each concentration is analyzed. A separate method blank shall be subjected to the analytical method along with the QC samples on each of the three (3) days. For each analyte, calculate the mean recovery for each day, for each level over days, and for all nine (9) samples. Calculate the relative standard deviation for each of the separate means obtained.
1484 | 1.5 Method Validation | 1.5.4 Measurement Uncertainty |  |  | V1M6 1.5.4 | Measurement Uncertainty
1485 | 1.5 Method Validation | 1.5.4 Measurement Uncertainty | All Measurements | Requirement | V1M6 1.5.4 | All radiochemical measurements shall provide the uncertainty of each quantitative measurement result.
1486 | 1.5 Method Validation | 1.5.4 Measurement Uncertainty | Comparison | Requirement | V1M6 1.5.4 | The results of the precision evaluation in Section 1.5.3 shall be compared to the uncertainty estimates as a check on the validity of the uncertainty evaluation procedures.
1487 | 1.5 Method Validation | 1.5.4 Measurement Uncertainty | Precision criterion | Requirement | V1M6 1.5.4 | The experimentally observed precision at each testing level shall not be statistically greater than the maximum combined standard uncertainty of the measurement results at that level, although it may be somewhat less.
1488 | 1.5 Method Validation | 1.5.4 Measurement Uncertainty | Combined Standard Uncertainty | Requirement | V1M6 1.5.4 | The combined standard uncertainty, when used, is the uncertainty of a measured value expressed as an estimated standard deviation.
1489 | 1.5 Method Validation | 1.5.4 Measurement Uncertainty | Combined Standard Uncertainty | Calculation | V1M6 1.5.4 | The combined standard uncertainty is calculated by combining the standard uncertainties of the input estimates.
1490 | 1.5 Method Validation | 1.5.5 Selectivity |  |  | V1M6 1.5.5 | Evaluation of Selectivity
1491 | 1.5 Method Validation | 1.5.5 Selectivity | Evaluation | Requirement | V1M6 1.5.5 | The laboratory shall evaluate selectivity, if applicable, by following the checks established within the method.
1492 | 1.6 DOC | 1.6.1 General |  |  | V1M6 1.6 | Demonstration of Capability (DOC)
1493 | 1.6 DOC | 1.6.1 General |  |  | V1M6 1.6.1 | General
1494 | 1.6 DOC | 1.6.1 General | Initial DOC | Requirement | V1M6 1.6.1 | Prior to acceptance and institution of any method for data reporting, satisfactory initial DOC is required (see Section 1.6.2).
1495 | 1.6 DOC | 1.6.1 General | Ongoing DOC | Requirement | V1M6 1.6.1 | Thereafter, ongoing DOC (Section 1.6.3), as per the quality control requirements in Section 1.7.3 (such as laboratory control samples) is required.
1496 | 1.6 DOC | 1.6.1 General | DOC | Exception | V1M6 1.6.1 | In cases where a laboratory analyzes samples using a method that has been in use by the laboratory for at least one year prior to applying for accreditation, and there have been no significant changes in instrument type, personnel or method, the ongoing DOC shall be acceptable as an initial DOC.
1497 | 1.6 DOC | 1.6.1 General | DOC | Record | V1M6 1.6.1 | The laboratory shall have records on file to demonstrate that an initial DOC is not required.
1498 | 1.6 DOC | 1.6.1 General | Documentation | Record | V1M6 1.6.1 | For the initial DOC, appropriate records as discussed in Section 1.6.2 shall be completed.
1499 | 1.6 DOC | 1.6.1 General | Frequency | Requirement | V1M6 1.6.1 | An initial DOC shall be completed each time there is a change in instrument type, personnel, or method (see 1.6.2).
1500 | 1.6 DOC | 1.6.1 General | Documentation | Requirement | V1M6 1.6.1 | All demonstrations shall be documented.
1501 | 1.6 DOC | 1.6.1 General | Data | Record | V1M6 1.6.1 | All data applicable to the demonstration shall be retained and readily available at the laboratory.
1502 | 1.6 DOC | 1.6.2 Initial DOC |  |  | V1M6 1.6.2 | Initial DOC
1503 | 1.6 DOC | 1.6.2 Initial DOC | DOC Frequency | Requirement | V1M6 1.6.2 | An initial DOC shall be made prior to using any method
1504 | 1.6 DOC | 1.6.2 Initial DOC | DOC Frequency | Requirement | V1M6 1.6.2 | An initial DOC shall be made at any time there is a change in instrument type, personnel or method
1505 | 1.6 DOC | 1.6.2 Initial DOC | DOC Frequency | Requirement | V1M6 1.6.2 | An initial DOC shall be made any time that a method has not been performed by the laboratory or analyst in a twelve (12) month period.
1506 | 1.6 DOC | 1.6.2 Initial DOC | Documentation | Record | V1M6 1.6.2.1 | The laboratory shall document each initial DOC in a manner such that the following information is readily available for each affected employee:
1507 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M6 1.6.2.1 a) | analyst(s) involved in preparation and/or analysis;
1508 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M6 1.6.2.1 b) | matrix;
1509 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M6 1.6.2.1 c) | analyte(s), class of analyte(s), or measured parameter(s);
1510 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M6 1.6.2.1 d) | identification of method(s) performed;
1511 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M6 1.6.2.1 e) | identification of laboratory-specific SOP used for analysis, including revision number;
1512 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M6 1.6.2.1 f) | date(s) of analysis;
1513 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M6 1.6.2.1 g) | summary of analyses, including information outlined in Section 1.6.2.2.c).
1514 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Requirement | V1M6 1.6.2.2 | If the method or regulation does not specify an initial DOC, the following procedure is acceptable. It is the responsibility of the laboratory to document that other approaches to initial DOC are adequate.
1515 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M6 1.6.2.2 a) | The analyte(s) shall be diluted in a volume of clean quality system matrix (a sample in which no target analytes or interferences are present at concentrations that will impact the results of a specific method) sufficient to prepare four (4) aliquots at a laboratory specified concentration. Where gamma-ray spectrometry is used to identify and quantify more than one analyte, the laboratory control sample shall contain gamma-emitting radionuclides that represent the low (e.g., 241Am), medium (e.g., 137Cs) and high (e.g., 60Co) energy range of the analyzed gamma-ray spectra. As indicated by these examples, the nuclides need not exactly bracket the calibrated energy range or the range over which nuclides are identified and quantified.
1516 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M6 1.6.2.2 b) | At least four (4) aliquots shall be prepared and analyzed according to the method either concurrently or over a period of days.
1517 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M6 1.6.2.2 c) | Using all of the results, calculate the mean recovery in the appropriate reporting units and the standard deviations of the population sample (in the same units) for each parameter of interest. When it is not possible to determine mean and standard deviations, such as for presence/absence and logarithmic values, the laboratory shall assess performance against established and documented criteria.
1518 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M6 1.6.2.2 d) | Compare the information from (c) above to the corresponding acceptance criteria for precision and accuracy in the method (if applicable) or in laboratory-generated acceptance criteria (if there are not established mandatory criteria). If all parameters meet the acceptance criteria, the analysis of actual samples may begin. If any one of the parameters does not meet the acceptance criteria, the performance is unacceptable for that parameter.
1519 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M6 1.6.2.2 e) | When one or more of the tested parameters fail at least one of the acceptance criteria, the analyst shall proceed according to i) or ii) below.
1520 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M6 1.6.2.2 e) i. | Locate and correct the source of the problem and repeat the test for all parameters of interest beginning with b) above.
1521 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M6 1.6.2.2 e) ii. | Beginning with b) above, repeat the test for all parameters that failed to meet criteria.
1522 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M6 1.6.2.2 f) | Repeated failure, however, confirms a general problem with the measurement system. If this occurs, locate and correct the source of the problem and repeat the test for all compounds of interest beginning with b).
1523 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Option | V1M6 1.6.2.2 g) | When an analyte not currently found on the laboratory’s list of accredited analytes is added to an existing accredited method, an initial DOC shall be performed for that analyte. When analytes are added to gamma-ray spectrometry and quantified this is not required.
1524 | 1.6 DOC | 1.6.3 Ongoing DOC |  |  | V1M6 1.6.3 | Ongoing DOC
1525 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Requirement | V1M6 1.6.3.1 | The laboratory shall have a documented procedure describing ongoing DOC.
1526 | 1.6 DOC | 1.6.3 Ongoing DOC | Competence | Requirement | V1M6 1.6.3.1 | The analyst(s) shall demonstrate ongoing capability by meeting the quality control requirements of the method, laboratory SOP, client specifications, and/or this Standard. It is the responsibility of the laboratory to document that other approaches to ongoing DOC are adequate.
1527 | 1.6 DOC | 1.6.3 Ongoing DOC |  |  | V1M6 1.6.3.2 | This on-going demonstration may include one of the following:
1528 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M6 1.6.3.2 a) | acceptable performance of a blind sample (single blind to the analyst);
1529 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M6 1.6.3.2 a) | Note: Successful analysis of a blind performance sample on a similar method using the same technology.
1530 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M6 1.6.3.2 b) | another initial DOC;
1531 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M6 1.6.3.2 c) | at least four (4) consecutive laboratory control samples with acceptable levels of precision and accuracy.
1532 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M6 1.6.3.2 c) | The laboratory shall determine the acceptable limits for precision and accuracy prior to analysis.
1533 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M6 1.6.3.2 c) | The laboratory shall tabulate or be able to readily retrieve four (4) consecutive passing LCS for each method for each analyst each year;
1534 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M6 1.6.3.2 d) | a documented process of analyst review using QC samples. QC samples can be reviewed to identify patterns for individuals or groups of analysts and determine if corrective action or retraining is necessary;
1535 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M6 1.6.3.2 e) | if a) through d) are not technically feasible, then analysis of real-world samples with results within predefined acceptance criteria (as defined by the laboratory or method) shall be performed.
1536 | 1.7 Technical Requirements |  |  |  | V1M6 1.7 | Technical Requirements
1537 | 1.7.1 Calibration |  |  |  | V1M6 1.7.1 | Instrument Calibration
1538 | 1.7.1 Calibration |  |  |  | V1M6 1.7.1 a) | Initial Calibration
1539 | 1.7.1 Calibration | Initial |  |  | V1M6 1.7.1 a) | This Section addresses those practices that are necessary for proper calibration of radiation counting instruments for environmental testing involving radioanalytical measurements.
1540 | 1.7.1 Calibration | Initial |  |  | V1M6 1.7.1 a) | This Section specifies the essential elements that shall define the procedures and documentation for initial instrument calibration and continuing instrument calibration verification to ensure that the data shall be of known quality and be appropriate for a given regulation or decision. This Standard does not specify detailed procedural steps (“how to”) for calibration, but establishes the essential elements for selection of the appropriate technique(s). This approach allows flexibility and permits the employment of a wide variety of analytical procedures and statistical approaches currently applicable for calibration.
1541 | 1.7.1 Calibration | Initial | Multiple/Conflicting Requirements | Requirement | V1M6 1.7.1 a) | If more stringent standards or requirements are included in a mandated method or regulation, the laboratory shall demonstrate that such requirements are met. If it is not apparent which standard is more stringent, then the requirements of the mandated method or regulation are to be followed.
1542 | 1.7.1 Calibration | Initial | Calibration | Exception | 1.7.1 a) | Given that radiation detection efficiency is essentially independent of sample activity at all but high activity levels (where dead time becomes significant), the requirements for calibration ranges of standards, of data reporting in calibration range, and the number of calibration standards are not applicable to radiochemical method calibrations except for mass attenuation in gas-proportional counting and sample quench in liquid scintillation counting.
1543 | 1.7.1 Calibration | Initial | Frequency | Requirement | V1M6 1.7.1 a) | Nuclear counting instruments are subject to calibration prior to initial use
1544 | 1.7.1 Calibration | Initial | Frequency | Requirement | V1M6 1.7.1 a) | Nuclear counting instruments are subject to calibration when the instrument is placed back into service after major repairs and the instrument’s response has changed as determined by a performance check
1545 | 1.7.1 Calibration | Initial | Frequency | Requirement | V1M6 1.7.1 a) | Nuclear counting instruments are subject to calibration when the instrument’s response exceeds predetermined acceptance criteria for the instrument quality control.
1546 | 1.7.1 Calibration | Initial | Frequency | Option | V1M6 1.7.1 a) | Instruments may also be recalibrated on a regular frequency even in the absence of these conditions.
1547 | 1.7.1 Calibration | Initial | Frequency | Procedure | V1M6 1.7.1 a) | The frequency of calibration shall be described in the laboratory method SOP if not specified in the method.
1548 | 1.7.1 Calibration | Initial | Frequency | Procedure | V1M6 1.7.1 a) | A specific frequency (e.g., annually) or calibrations based on observations from the associated control or tolerance chart, shall be specified in the laboratory method SOP.
1549 | 1.7.1 Calibration | Initial | Reference Standard | Requirement | V1M6 1.7.1 a) | Instrument calibration shall be performed with reference standards as defined in Section 1.7.2.5.c).
1550 | 1.7.1 Calibration | Initial | Standard Characteristics | Requirement | V1M6 1.7.1 a) | The standards shall have the same general characteristics (i.e., geometry, homogeneity, density, etc.) as the associated samples.
1551 | 1.7.1 Calibration | Initial |  |  | V1M6 1.7.1 a) | The following items are essential elements of initial instrument calibration:
1552 | 1.7.1 Calibration | Initial | Processing | Procedure | V1M6 1.7.1 a) i. | The details of the initial instrument calibration procedures including calculations, acceptance criteria and associated statistics shall be included or referenced in the method SOP.
1553 | 1.7.1 Calibration | Initial | Calibration Procedure | Record | V1M6 1.7.1 a) i. | When initial instrument calibration procedures are referenced in the method, then the referenced material shall be retained by the laboratory and be available for review.
1554 | 1.7.1 Calibration | Initial | Data | Record | V1M6 1.7.1 a) ii. | Sufficient raw data records shall be retained to permit reconstruction of the initial instrument calibration (e.g., calibration date, method, instrument, analysis date, each analyte name, analyst’s initials or signature;
1555 | 1.7.1 Calibration | Initial | Quantitation | Requirement | V1M6 1.7.1 a) iii. | Sample results shall be quantitated from the initial instrument calibration
1556 | 1.7.1 Calibration | Initial | Quantitation | Exception | V1M6 1.7.1 a) iii. | Sample results may not be quantitated from any continuing instrument calibration verification unless otherwise required by regulation, method, or program.
1557 | 1.7.1 Calibration | Initial | Second Source Verification | Requirement | V1M6 1.7.1 a) iv. | All initial instrument calibrations shall be verified with a standard obtained from a second manufacturer or lot if the lot can be demonstrated from the manufacturer as prepared independently from other lots.
1558 | 1.7.1 Calibration | Initial | Second Source Standard Traceability | Requirement | V1M6 1.7.1 a) iv. | Traceability shall be to a national standard, when commercially available.
1559 | 1.7.1 Calibration | Initial | Acceptance Criteria | Procedure | V1M6 1.7.1 a) v. | Criteria for the acceptance of an initial instrument calibration shall be established (e.g., correlation coefficient or relative percent difference).
1560 | 1.7.1 Calibration | Initial | Applicability | Requirement | V1M6 1.7.1 a) v. | The criteria used shall be appropriate to the calibration technique employed.
1561 | 1.7.1 Calibration | Initial | Unacceptable Calibration | Corrective Action | V1M6 1.7.1 a) vi. | If the initial instrument calibration results are outside established acceptance criteria, corrective actions shall be performed
1562 | 1.7.1 Calibration | Initial | Unacceptable Calibration | Requirement | V1M6 1.7.1 a) vi. | If the initial instrument calibration results are outside established acceptance criteria all associated samples re-analyzed.
1563 | 1.7.1 Calibration | Initial | Unacceptable Calibration | Reporting Requirement | V1M6 1.7.1 a) vi. | If re-analysis of the samples is not possible, data associated with an unacceptable initial instrument calibration shall be reported with appropriate data qualifiers.
1564 | 1.7.1 Calibration | Initial | Number of Standards | Procedure | V1M6 1.7.1 a) vii. | If a reference or mandated method does not specify the number of calibration standards, the laboratory shall have a written procedure for determining the number of points for establishing the initial instrument calibration.
1565 | 1.7.1 Calibration | Calibration Verification |  |  | V1M6 1.7.1 b) | Instrument Calibration Verification (Performance Checks)
1566 | 1.7.1 Calibration | Calibration Verification | Monitoring and Evaluation | Requirement | V1M6 1.7.1 b) | Performance checks shall be performed using appropriate check sources and monitored with control charts or tolerance charts to ensure that the instrument is operating properly, the detector response has not significantly changed, and therefore the instrument calibration has not changed.
1567 | 1.7.1 Calibration | Calibration Verification | Tolerance Charts | Requirement | V1M6 1.7.1 b) | The same check source used in the preparation of the tolerance chart or control chart at the time of calibration shall be used in the calibration verification of the instrument (performance checks).
1568 | 1.7.1 Calibration | Calibration Verification | Check Sources | Specifications | V1M6 1.7.1 b) | The check sources shall provide adequate counting statistics for a relatively short count time and the source should be sealed or encapsulated to prevent loss of activity and contamination of the instrument and laboratory personnel.
1569 | 1.7.1 Calibration | Calibration Verification | Gamma Ray | Frequency | V1M6 1.7.1 b) i. | For gamma-ray spectroscopy systems, performance checks for detection efficiency shall be performed on a day-of-use basis.
1570 | 1.7.1 Calibration | Calibration Verification | Gamma Ray | Frequency | V1M6 1.7.1 b) i. | For gamma-ray spectroscopy systems, performance checks for energy calibration shall be performed on a day-of-use basis.
1571 | 1.7.1 Calibration | Calibration Verification | Gamma Ray | Frequency | V1M6 1.7.1 b) i. | For gamma-ray spectroscopy systems, performance checks for peak resolution shall be performed on a day-of-use basis.
1572 | 1.7.1 Calibration | Calibration Verification | Alpha Particle | Frequency | V1M6 1.7.1 b) ii. | For alpha-particle spectroscopy systems, the performance check for energy calibration shall be performed on a weekly basis,
1573 | 1.7.1 Calibration | Calibration Verification | Alpha Particle | Frequency | V1M6 1.7.1 b) ii. | and the performance check for detection efficiency shall be performed on at least a monthly basis.
1574 | 1.7.1 Calibration | Calibration Verification | Gas Proportional and Liquid Scintillation Counters | Frequency | V1M6 1.7.1 b) iii. | For gas-proportional and liquid scintillation counters, the performance check for detection efficiency shall be performed on a day-of-use basis.
1575 | 1.7.1 Calibration | Calibration Verification | Gas Proportional and Liquid Scintillation Counters | Frequency | V1M6 1.7.1 b) iii. | For batches of samples that uninterruptedly count for more than a day, a performance check may be performed instead at the beginning and end of the batch as long as this time interval is no greater than one week.
1576 | 1.7.1 Calibration | Calibration Verification | Scintillation Counters | Frequency | V1M6 1.7.1 b) iv. | For scintillation counters the calibration verification for detection efficiency shall be performed on a day-of-use basis.
1577 | 1.7.1 Calibration | Background Measurement |  |  | V1M6 1.7.1 c) | Background Measurement
1578 | 1.7.1 Calibration | Background Measurement | Control chart Evaluation | Requirement | V1M6 1.7.1 c) | Background measurements shall be made on a regular basis and monitored using control charts or tolerance charts to ensure that a laboratory maintains its capability to meet required measurement quality objectives. This background measurement is not the short term check for contamination that is addressed in 1.7.1 d).
1579 | 1.7.1 Calibration | Background Measurement | Use | Requirement | V1M6 1.7.1 c) | The background measurement values must be subtracted from the total measured activity in the determination of the sample activity.
1580 | 1.7.1 Calibration | Background Measurement | Gamma Ray | Frequency | V1M6 1.7.1 c) i. | For gamma-ray spectroscopy systems, background measurements shall be performed on at least a monthly basis.
1581 | 1.7.1 Calibration | Background Measurement | Alpha Particle | Frequency | V1M6 1.7.1 c) ii. | For alpha-particle spectroscopy systems, background measurements shall be performed on at least a monthly basis.
1582 | 1.7.1 Calibration | Background Measurement | Gas Proportional | Frequency | V1M6 1.7.1 c) iii. | For gas-proportional counters background measurements shall be performed on at least a weekly basis.
1583 | 1.7.1 Calibration | Background Measurement | Scintillation Counters | Frequency | V1M6 1.7.1 c) iv. | For scintillation counters, background measurements shall be performed each day of use.
1584 | 1.7.1 Calibration | Instrument Contamination Monitoring |  |  | V1M6 1.7.1 d) | Instrument Contamination Monitoring
1585 | 1.7.1 Calibration | Instrument Contamination Monitoring | Monitoring for contamination | Procedure | V1M6 1.7.1 d) | The laboratory shall have a written procedure for monitoring radiation measurement instrumentation for radioactive contamination.
1586 | 1.7.1 Calibration | Instrument Contamination Monitoring | Procedure | Details | V1M6 1.7.1 d) | The procedure shall indicate the frequency of the monitoring and shall indicate criteria, which initiates corrective action.
1587 | 1.7.2 QC |  |  |  | V1M6 1.7.2 | Quality Control for Radiochemistry
1588 | 1.7.2 QC | General |  | Procedure | V1M6 1.7.2 | The laboratory shall have quality control procedures for monitoring the validity of environmental tests undertaken as specified in this Section. (see requirements below)
1589 | 1.7.2 QC | General | Monitoring | Requirement | V1M6 1.7.2 | This monitoring shall be planned and reviewed.
1590 | 1.7.2 QC | General | Corrective Action | Requirement | V1M6 1.7.2 | The failure of any QC sample analysis and the corrective actions taken shall be noted in the laboratory report for the associated samples.
1591 | 1.7.2 QC | 1.7.2.1 Negative Control | Method Blank |  | V1M6 1.7.2.1 | Negative Control – Method Performance: Method Blank
1592 | 1.7.2 QC | 1.7.2.1 Negative Control | Purpose | Requirement | V1M6 1.7.2.1 a) | The method blank is used to assess the preparation batch for possible contamination during the preparation and processing steps or for other low-level bias.
1593 | 1.7.2 QC | 1.7.2.1 Negative Control | Processing | Requirement | V1M6 1.7.2.1 a) | The method blank shall be processed along with and under the same conditions as the associated samples to include all steps of the analytical procedure.
1594 | 1.7.2 QC | 1.7.2.1 Negative Control | Blank Assessment | Procedure | V1M6 1.7.2.1 a) | Procedures shall be in place to determine if a method blank result is significantly different from zero.
1595 | 1.7.2 QC | 1.7.2.1 Negative Control | Reporting | Requirement | V1M6 1.7.2.1 a) | Any affected samples associated with a failed method blank shall be reprocessed for analysis or the results reported with appropriate data-qualifying codes.
1596 | 1.7.2 QC | 1.7.2.1 Negative Control | Frequency | Requirement | V1M6 1.7.2.1 b) | The method blank shall be analyzed at a minimum of one (1) per preparation batch, which shall be a maximum of twenty (20) field samples, for all radiochemical methods except gross alpha/beta in solid matrices and gamma-ray spectrometry.
1597 | 1.7.2 QC | 1.7.2.1 Negative Control | Matrix Attributes | Requirement | V1M6 1.7.2.1 c) | The method blank shall consist of a quality system matrix that is similar to the associated samples and is known to be as free of the analytes of interest as possible.
1598 | 1.7.2 QC | 1.7.2.1 Negative Control | Blank Subtraction | Requirement | V1M6 1.7.2.1 c) | There shall be no subtraction of the method blank result from the sample results in the associated preparation or analytical batch unless permitted by method or program.
1599 | 1.7.2 QC | 1.7.2.1 Negative Control | Blank Subtraction | Exception | V1M6 1.7.2.1 c) | This requirement does not preclude corrections for background radiation (e.g., instrument background, analyte in the tracer or carrier, reagent impurities, peak overlap, etc.) to all analyzed samples, both program/project submitted and internal quality control samples. However, these corrections shall not depend on the result of the method blank analysis, whose purpose is to check for uncorrected contamination or other low-level bias.
1600 | 1.7.2 QC | 1.7.2.1 Negative Control | Blank Preparation | Requirement | V1M6 1.7.2.1 c) | The method blank sample shall be prepared with aliquot size similar to that of the routine samples for analysis.
1601 | 1.7.2 QC | 1.7.2.2 Positive Control | LCS |  | V1M6 1.7.2.2 | Positive Control – Method Performance: Laboratory Control Sample (LCS)
1602 | 1.7.2 QC | 1.7.2.2 Positive Control | Purpose | Requirement | V1M6 1.7.2.2 a) | The LCS is used to evaluate the performance of the total analytical system, including all preparation and analysis steps.
1603 | 1.7.2 QC | 1.7.2.2 Positive Control | Evaluation | Requirement | V1M6 1.7.2.2 a) | Results of the LCS are compared to established criteria and, if found to be outside of these criteria may indicate that the analytical system is “out of control.”
1604 | 1.7.2 QC | 1.7.2.2 Positive Control | Unacceptable LCS | Requirement | V1M6 1.7.2.2 a) | Any affected samples associated with an out-of-control LCS shall be reprocessed for reanalysis or the results reported with appropriate data qualifying codes.
1605 | 1.7.2 QC | 1.7.2.2 Positive Control | Frequency | Requirement | V1M6 1.7.2.2 b) | The LCS shall be analyzed at a minimum of one per preparation batch. Exceptions would be for those analytes for which no spiking solutions are available.
1606 | 1.7.2 QC | 1.7.2.2 Positive Control | Matrix Attributes | Requirement | V1M6 1.7.2.2 c) | The LCS is a quality system matrix, known to be free of analytes of interest, spiked with known and verified concentrations of analytes.
1607 | 1.7.2 QC | 1.7.2.2 Positive Control | Matrix Attributes | Option | V1M6 1.7.2.2 c) | NOTE: The matrix spike may be used in place of this control as long as the acceptance criteria are as stringent as for the LCS.
1608 | 1.7.2 QC | 1.7.2.2 Positive Control | Matrix Attributes | Option | V1M6 1.7.2.2 d) | Alternatively the LCS may consist of a medium containing known and verified concentrations of analytes or as Certified Reference Material (CRM). The components to be spiked shall be as specified by the mandated method or regulation or as requested by the client.
1609 | 1.7.2 QC | 1.7.2.2 Positive Control | Activity | Requirement | V1M6 1.7.2.2 e) | The activity of the laboratory control sample shall be: (1) at least ten (10) times the MDA, and (2) at a level comparable to that of routine samples when such information is available if the sample activities are expected to exceed ten times the MDA.
1610 | 1.7.2 QC | 1.7.2.2 Positive Control | LCS Source | Requirement | V1M6 1.7.2.2 f) | The laboratory standards used to prepare the laboratory control sample shall be from a source independent of the laboratory standards used for instrument calibration,
1611 | 1.7.2 QC | 1.7.2.2 Positive Control | Reference Standard Criteria | Requirement | V1M6 1.7.2.2 f) | The laboratory standards used to prepare the laboratory control sample shall meet the requirements for reference standards provided in Section 1.7.5.2.c).
1612 | 1.7.2 QC | 1.7.2.2 Positive Control | Multiple Isotope Methods | Requirement | V1M6 1.7.2.2 g) | Where a radiochemical method, other than gamma-ray spectroscopy, has more than one reportable analyte isotope (e.g. plutonium, 238Pu and 239Pu, using alpha-particle spectrometry), only one of the analyte isotopes need be included in the laboratory control sample at the indicated activity level.
1613 | 1.7.2 QC | 1.7.2.2 Positive Control | Multiple Isotope Methods | Requirement | V1M6 1.7.2.2 g) | However, where more than one analyte is detectable, each shall be assessed against the specified acceptance criteria.
1614 | 1.7.2 QC | 1.7.2.2 Positive Control | Gamma Ray Energy Ranges | Requirement | V1M6 1.7.2.2 h) | Where gamma-ray spectrometry is used to identify and quantify more than one analyte, the laboratory control sample shall contain gamma-emitting radionuclides that represent the low (e.g., 241Am), medium (e.g., 137Cs) and high (e.g., 60Co) energy range of the analyzed gamma-ray spectra. As indicated by these examples, the nuclides need not exactly bracket the calibrated energy range or the range over which nuclides are identified and quantified.
1615 | 1.7.2 QC | 1.7.2.2 Positive Control | Blank Preparation | Requirement | V1M6 1.7.2.2 i) | The laboratory control sample shall be prepared with similar aliquot size to that of the routine samples for analyses.
1616 | 1.7.2 QC | 1.7.2.3 Sample-Specific Controls |  |  | V1M6 1.7.2.3 | Sample-Specific Controls
1617 | 1.7.2 QC | 1.7.2.3 Sample-Specific Controls | Matrix Effects | Procedure | V1M6 1.7.2.3 | The laboratory shall document procedures for determining the effect of the sample matrix on method performance.
1618 | 1.7.2 QC | 1.7.2.3 Sample-Specific Controls | Matrix Effects | Procedure | V1M6 1.7.2.3 | These procedures relate to the analyses of quality system matrix specific quality control (QC) samples and are designed as data quality indicators for a specific sample using the designated method.
1619 | 1.7.2 QC | 1.7.2.3 Sample-Specific Controls | Matrix Tracking | Procedure | V1M6 1.7.2.3 | Examples of matrix-specific QC include: Matrix Spike (MS); Matrix Spike Duplicate (MSD); and replicates. The laboratory shall have procedures in place for tracking, managing, and handling matrix-specific QC criteria including spiking appropriate components at appropriate concentrations, calculating recoveries and relative percent difference, evaluating and reporting results based on performance of the QC samples.
1620 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes |  |  | V1M6 1.7.2.3 a) | Matrix Spike
1621 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes |  |  | V1M6 1.7.2.3 a) i. | Matrix spikes indicate the effect of the sample matrix on the accuracy of the results generated using the selected method. The results of this analysis shall be one of the quality control measures used to assess the batch.
1622 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes | Frequency | Requirement | V1M6 1.7.2.3 a) ii. | The frequency of the analysis of matrix spikes are as specified by the method or may be determined as part of the contract review process.
1623 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes | Spiked Components | Requirement | V1M6 1.7.2.3 a) iii. | The components to be spiked shall be as specified by the mandated method.
1624 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes | Spiked Components | Requirement | V1M6 1.7.2.3 a) iii. | Any permit specified analytes, as specified by regulation or client requested analytes shall also be included.
1625 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes | Insufficient Sample | Requirement | V1M6 1.7.2.3 a) iv. | The lack of sufficient sample aliquot size to perform a matrix spike shall be noted in the laboratory report.
1626 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes | Activity | Requirement | V1M6 1.7.2.3 a) v. | The activity of the matrix spike analytes(s) shall be greater than five times the MDA.
1627 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes | Second Source | Requirement | V1M6 1.7.2.3 a) vi. | The laboratory standards used to prepare the matrix spike shall be from a source independent of the laboratory standards used for instrument calibration,
1628 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes | Reference Standard Criteria | Requirement | V1M6 1.7.2.3 a) vi. | The laboratory standards used to prepare the matrix spike shall meet the requirements for reference standards of Section 1.7.2.5.c).
1629 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes | Preparation | Requirement | V1M6 1.7.2.3 a) vii. | The matrix spike shall be prepared by adding a known activity of target analyte after sub-sampling if required but before any chemical treatment (e.g., chemical digestion, dissolution, separation, etc.).
1630 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes | Multiple Isotope Methods | Requirement | V1M6 1.7.2.3 a) vii. | Where a radiochemical method, other than gamma-ray spectroscopy, has more than one reportable analyte isotope (e.g. plutonium, 238Pu and 239Pu, using alpha-particle spectrometry), only one of the analyte isotopes need be included in the matrix spike sample at the indicated activity level.
1631 | 1.7.2 QC | 1.7.2.3 a) Matrix Spikes | Multiple Isotope Methods | Requirement | V1M6 1.7.2.3 a) vii. | However, where more than one analyte is detectable, each shall be assessed against the specified acceptance criteria.
1632 | 1.7.2 QC | 1.7.2.3 b) Replicates |  |  | V1M6 1.7.2.3 b) | Replicates / Matrix Spike Duplicates / Laboratory Control Sample Duplicates
1633 | 1.7.2 QC | 1.7.2.3 b) Replicates | Preparation | Requirement | V1M6 1.7.2.3 b) i. | Replicates are defined as replicate aliquots of the same sample taken through the entire analytical procedure.
1634 | 1.7.2 QC | 1.7.2.3 b) Replicates | Intended Use | Requirement | V1M6 1.7.2.3 b) i. | The results from replicate analysis indicate the precision of the results for the specific sample using the selected method.
1635 | 1.7.2 QC | 1.7.2.3 b) Replicates | Target Analytes | Occurrence | V1M6 1.7.2.3 b) i. | Replicates provide the most useful measure of precision when target analytes are found in the sample chosen for replication.
1636 | 1.7.2 QC | 1.7.2.3 b) Replicates | Frequency | Requirement | V1M6 1.7.2.3 b) ii. | The frequency of the analysis of matrix replicates and duplicates are as specified by the method or may be determined as part of the contract review process.
1637 | 1.7.2 QC | 1.7.2.3 b) Replicates | Sample Source | Requirement | V1M6 1.7.2.3 b) iii. | Replicates are performed on replicate aliquots of actual samples.
1638 | 1.7.2 QC | 1.7.2.3 b) Replicates | Low-level | Option | V1M6 1.7.2.3 b) iv. | For low-level samples (less than approximately three times the MDA) the laboratory may analyze a laboratory control samples duplicate or a replicate matrix spike (matrix spike and a matrix spike duplicate) to determine reproducibility within a preparation batch in place of a sample replicate.
1639 | 1.7.2 QC | 1.7.2.3 b) Replicates | LCS Duplicate | Option | V1M6 1.7.2.3 b) iv. | In addition based on project or program requirements, the laboratory may analyze a laboratory control sample duplicate or a matrix spike duplicate in place of a sample replicate.
1640 | 1.7.2 QC | 1.7.2.3 c) Tracer |  |  | V1M6 1.7.2.3 c) | Tracer
1641 | 1.7.2 QC | 1.7.2.3 c) Tracer | Calculation & Reporting | Requirement | V1M6 1.7.2.3 c) | For those methods that employ a tracer for yield determination, each sample result shall have an associated tracer yield calculated and reported.
1642 | 1.7.2 QC | 1.7.2.3 c) Tracer | Processing | Requirement | V1M6 1.7.2.3 c) | The tracer shall be added to the sample after subsampling, if required, but before any chemical treatment (e.g., chemical digestion, dissolution, separation, etc.) unless otherwise specified by the method.
1643 | 1.7.2 QC | 1.7.2.3 c) Tracer | Tracer yield | Requirement | V1M6 1.7.2.3 c) | The tracer yield for each sample result shall be one of the quality control measures to be used to assess the associated sample result acceptance.
1644 | 1.7.2 QC | 1.7.2.3 c) Tracer | Assessment | Requirement | V1M6 1.7.2.3 c) | The tracer yield shall be assessed against the specific acceptance criteria specified in the method.
1645 | 1.7.2 QC | 1.7.2.3 c) Tracer | Corrective Action | Requirement | V1M6 1.7.2.3 c) | When the specified tracer yield acceptance criteria are not met, the specified corrective action and contingencies shall be followed.
1646 | 1.7.2 QC | 1.7.2.3 c) Tracer | Reporting | Requirement | V1M6 1.7.2.3 c) | The occurrence of a failed tracer yield and the actions taken shall be noted in the laboratory report to the client.
1647 | 1.7.2 QC | 1.7.2.3 d) Carrier |  |  | V1M6 1.7.2.3 d) | Carrier
1648 | 1.7.2 QC | 1.7.2.3 d) Carrier | Calculation & Reporting | Requirement | V1M6 1.7.2.3 d) | For those methods that utilize a carrier for yield determination, each sample shall have an associated carrier yield calculated and reported.
1649 | 1.7.2 QC | 1.7.2.3 d) Carrier | Processing | Requirement | V1M6 1.7.2.3 d) | The carrier shall be added to the sample after subsampling, if required, but before any chemical treatment (e.g., chemical digestion, dissolution, separation, etc.) unless otherwise specified by the method.
1650 | 1.7.2 QC | 1.7.2.3 d) Carrier | Tracer yield | Requirement | V1M6 1.7.2.3 d) | The carrier yield for each sample shall be one of the quality control measures to be used to assess the associated sample result acceptance.
1651 | 1.7.2 QC | 1.7.2.3 d) Carrier | Assessment | Requirement | V1M6 1.7.2.3 d) | The carrier yield shall be assessed against the specific acceptance criteria specified in the laboratory method SOP.
1652 | 1.7.2 QC | 1.7.2.3 d) Carrier | Corrective Action | Requirement | V1M6 1.7.2.3 d) | When the specified carrier yield acceptance criteria are not met, the specified corrective action and contingencies shall be followed.
1653 | 1.7.2 QC | 1.7.2.3 d) Carrier | Reporting | Requirement | V1M6 1.7.2.3 d) | The occurrence of a failed carrier yield and the actions taken shall be noted in the laboratory report to the client.
1654 | 1.7.2 QC | 1.7.2.4 Data Reduction |  |  | V1M6 1.7.2.4 | Data Reduction
1655 | 1.7.2 QC | 1.7.2.4 Data Reduction | Data Reduction | Procedure | V1M6 1.7.2.4 a) | The procedures for data reduction, such as use of linear regression, shall be documented.
1656 | 1.7.2 QC | 1.7.2.4 Data Reduction | Measurement Uncertainty | Requirement | V1M6 1.7.2.4 b) | Measurement Uncertainties. Each result shall be reported with its measurement uncertainty. The report should clearly explain the uncertainty. At a minimum the report shall:
1657 | 1.7.2 QC | 1.7.2.4 Data Reduction | Uncertainty Type | Requirement | V1M6 1.7.2.4 b) i. | indicate whether the uncertainty is the combined standard uncertainty (“one sigma”) or an expanded uncertainty; and
1658 | 1.7.2 QC | 1.7.2.4 Data Reduction | Expanded Uncertainty | Requirement | V1M6 1.7.2.4 b) ii. | for expanded uncertainties, indicate the coverage factor (k) and optionally the approximate level of confidence.
1659 | 1.7.2 QC | 1.7.2.4 Data Reduction | Procedure Reference | Requirement | V1M6 1.7.2.4 c) | The procedures for determining the measurement uncertainty shall be documented and shall be consistent with the ISO Guide 98: 1995, Guide to the Expression of Uncertainty in Measurement (GUM) and with the recommendations of Chapter 19 of the Multi-Agency Radiological Laboratory Analytical Protocols Manual (MARLAP) Volume I (EPA 402-B-04-001A), Volume II (EPA 402-B-04-001B), Volume III (EPA 402-B-04-001C), July 2004.
1660 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks |  |  | V1M6 1.7.2.5 | 1.7.2.5 Reagent Quality, Water Quality, and Checks
1661 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Purity | Requirement | V1M6 1.7.2.5 a) | In methods where the purity of reagents is not specified, reagents shall be analytical reagent grade or better. Reagents of lesser purity than those specified by the method shall not be used.
1662 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Verification | Requirement | V1M6 1.7.2.5 a) | The labels on the container should be checked to verify that the purity of the reagents meets the requirements of the particular method. Such information shall be available.
1663 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Water Source Quality | Requirement | V1M6 1.7.2.5 b) | The quality of water sources shall be monitored
1664 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Water Source Quality | Requirement | V1M6 1.7.2.5 b) | The quality of water sources shall be documented
1665 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Water Source Quality | Requirement | V1M6 1.7.2.5 b) | The quality of water sources shall meet method specified requirements.
1666 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Radionuclide Standards | Procedure | V1M6 1.7.2.5 c) | The quality control program shall establish and maintain provisions for radionuclide standards. (1.7.2.5 c) iii.)
1667 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Source | Requirement | V1M6 1.7.2.5 c) i. | Reference standards that are used in a radiochemical laboratory shall be obtained from NIST or suppliers who participate in supplying NIST standards or NIST traceable radionuclides.
1668 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Source | Requirement | V1M6 1.7.2.5 c) i. | Any reference standards purchased outside the United States shall be traceable back to each country's national standards laboratory.
1669 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Commercial Suppliers | Requirement | V1M6 1.7.2.5 c) i. | Commercial suppliers of reference standards shall conform to ANSI N42.22 to assure the quality of their products.
1670 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Certificate of Calibration | Requirement | V1M6 1.7.2.5 c) ii. | Reference standards shall be accompanied with a certificate of calibration whose content is as described in ANSI N42.22 - 1995, Section 8, Certificates.
1671 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Deviations from value | Requirement | V1M6 1.7.2.5 c) iii. | Laboratories should consult with the supplier if the lab's verification of the activity of the reference traceable standard indicates a noticeable deviation from the certified value.
1672 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Value to use | Requirement | V1M6 1.7.2.5 c) iii. | The laboratory shall use only the decay-corrected certified value.
1673 | 1.7.2 QC | 1.7.2.5 Reagent Quality, Water Quality, and Checks | Handling, storage and expiration | Procedure | V1M6 1.7.2.5 c) iii. | The laboratory shall have a written procedure for handling, storing, and establishing expiration dates for reference standards.
1674 | 1.7.2 QC | 1.7.2.6 Selectivity |  |  | V1M6 1.7.2.6 | Selectivity
1675 | 1.7.2 QC | 1.7.2.6 Selectivity | Evaluation | Requirement | V1M6 | The laboratory shall evaluate selectivity by following the checks established within the method.
1676 | 1.7.2 QC | 1.7.2.7 Constant and Consistent Test Conditions |  |  | V1M6 1.7.2.7 | Constant and Consistent Test Conditions
1677 | 1.7.2 QC | 1.7.2.7 Constant and Consistent Test Conditions | Consistent Operation | Requirement | V1M6 1.7.2.7 a) | The laboratory shall assure that the test instruments consistently operate within the specifications required of the application for which the equipment is used.
1678 | 1.7.2 QC | 1.7.2.7 Constant and Consistent Test Conditions | Glassware | Requirement | V1M6 1.7.2.7 b) | Glassware Cleaning. Glassware shall be cleaned to meet the sensitivity requirements of the method.
1679 | 1.7.2 QC | 1.7.2.7 Constant and Consistent Test Conditions | Glassware | Procedure | V1M6 1.7.2.7 b) | Any cleaning and storage procedures that are not specified by the method shall be documented in laboratory records and SOPs. Note that some applications may require single-use glassware.
1680 | 1.7.2 QC | 1.7.2.7 Constant and Consistent Test Conditions | Radiological Control Program | Requirement | V1M6 1.7.2.7 c) | Radiological Control Program. The laboratory shall maintain a radiological control program that addresses analytical radiological control.
1681 | 1.7.2 QC | 1.7.2.7 Constant and Consistent Test Conditions | Radiological Control Program | Procedure | V1M6 1.7.2.7 c) | The program shall address the procedures for segregating samples with potentially widely varying levels of radioactivity.
1682 | 1.7.2 QC | 1.7.2.7 Constant and Consistent Test Conditions | Radiological Control Program | Procedure | V1M6 1.7.2.7 c) | The radiological control program shall explicitly define how low-level and high-level samples will be identified, segregated and processed in order to prevent sample cross-contamination.
1683 | 1.7.2 QC | 1.7.2.7 Constant and Consistent Test Conditions | Radiological Control Program | Procedure | V1M6 1.7.2.7 c) | The radiological control program shall include the measures taken to monitor and evaluate background activity or contamination on an ongoing basis.
1684 | 1.7.3 Data Acceptance / Rejection Criteria |  |  |  | V1M6 | Data Acceptance/Rejection Criteria
1685 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.1 Negative Control |  |  | V1M6 | Negative Control – Method Performance: Method Blank
1686 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.1 Negative Control | Evaluation | Requirement | V1M6 1.7.3.1 a) | While the goal is to have no statistically significant difference from zero, each method blank shall be critically evaluated as to the nature of the interference and the effect on the analysis of each sample within the batch.
1687 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.1 Negative Control | Contamination Source | Requirement | V1M6 1.7.3.1 a) | The source of contamination or other bias shall be investigated and measures taken to minimize or eliminate the problem and affected samples reprocessed or data shall be appropriately qualified if:
1688 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.1 Negative Control | Contamination | Exception | V1M6 1.7.3.1 a) i. | the absolute value of the activity of a targeted analyte in the blank exceeds three times its combined standard uncertainty, AND is greater than 1/10 of the activity measured in any sample; or
1689 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.1 Negative Control | Contamination | Exception | V1M6 1.7.3.1 a) ii. | the method blank result otherwise affects the sample results as per the method requirements or the project-specific measurement quality objectives.
1690 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.1 Negative Control | Aliquot size compensation | Requirement | V1M6 1.7.3.1 b) | The acceptance criteria for samples associated with a failed method blank shall be calculated in a manner that compensates for sample results based on differing aliquot sizes.
1691 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.1 Negative Control | Investigation | Requirement | V1M6 1.7.3.1 c) | When a blank result is determined to be significantly different from zero, the cause shall be investigated and measures taken to minimize or eliminate the problem.
1692 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.1 Negative Control | Corrective Action | Requirement | V1M6 1.7.3.1 c) | Samples associated with a failed blank shall be evaluated as to the best corrective action for the samples (e.g., reprocessing or data qualifying codes).
1693 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.1 Negative Control | Reporting | Requirement | V1M6 1.7.3.1 d) | The occurrence of a failed method blank and any associated corrective action shall be noted in the laboratory report to the client.
1694 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.2 Positive Control |  |  | V1M6 1.7.3.2 | Positive Control – Method Performance: Laboratory Control Sample (LCS)
1695 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.2 Positive Control | Evaluation | Requirement | V1M6 1.7.3.2 a) | a) The results of the individual batch LCS are calculated in percent recovery or other appropriate statistical technique that allows comparison to established acceptance criteria. (1.7.3.2.b)
1696 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.2 Positive Control | Calculation | Requirement | V1M6 1.7.3.2 a) | The laboratory shall document the calculation.
1697 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.2 Positive Control | Evaluation | Requirement | V1M6 1.7.3.2 b) | The individual LCS is compared to the acceptance criteria as published in the mandated method.
1698 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.2 Positive Control | Evaluation | Requirement | V1M6 1.7.3.2 b) | Where there are no established criteria, the laboratory shall determine internal criteria and document the method used to establish the limits or utilize client specified assessment criteria.
1699 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.2 Positive Control | Acceptable Performance |  | V1M6 1.7.3.2 c) | An LCS that is determined to be within the criteria effectively establishes that the analytical system is in control and validates system performance for the samples in the associated batch.
1700 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.2 Positive Control | Unacceptable Performance | Requirement | V1M6 1.7.3.2 c) | Samples analyzed along with an LCS determined to be “out of control” shall be considered suspect and the samples reprocessed and re-analyzed or the data reported with appropriate data qualifying codes.
1701 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.2 Positive Control | Reporting | Requirement | V1M6 1.7.3.2 d) | The occurrence of a failed LCS and any associated actions shall be noted in the laboratory report to the client.
1702 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 Sample-Specific Controls |  |  | V1M6 1.7.3.3 | Sample-Specific Controls
1703 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 a) MS/MSDs |  |  | V1M6 1.7.3.3 a) | Matrix Spike; Matrix Spike Duplicates
1704 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 a) MS/MSDs |  |  | V1M6 1.7.3.3 a) i. | The results from matrix spike/matrix spike duplicate are primarily designed to assess the precision and accuracy of analytical results in a given matrix and are expressed as percent recovery (%R), relative percent difference (RPD), or other appropriate statistical technique that allows comparison to established acceptance criteria.
1705 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 a) MS/MSDs | Calculation | Requirement | V1M6 1.7.3.3 a) i. | The laboratory shall document the calculation for %R, RPD or other statistical treatment used.
1706 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 a) MS/MSDs | Evaluation | Requirement | V1M6 1.7.3.3 a) ii. | The results are compared to the acceptance criteria as published in the mandated test method.
1707 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 a) MS/MSDs | Evaluation | Requirement | V1M6 1.7.3.3 a) ii. | Where there are no established criteria, the laboratory shall determine internal criteria and document the method used to establish the limits.
1708 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 a) MS/MSDs | Corrective Action | Requirement | V1M6 1.7.3.3 a) ii. | For matrix spike results outside established criteria, corrective action shall be documented or the data reported with appropriate data qualifying codes.
1709 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 a) MS/MSDs | Reporting | Requirement | V1M6 1.7.3.3 a) iii. | The occurrence of a failed matrix spike and any associated actions shall be noted in the laboratory report to the client.
1710 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 b) Replicates |  |  | V1M6 1.7.3.3 b) | Replicates
1711 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 b) Replicates |  |  | V1M6 1.7.3.3 b) i. | The results from replicates are primarily designed to assess the precision of analytical results in a given matrix and are expressed as relative percent difference (RPD) or another statistical treatment (e.g., normalized differences).
1712 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 b) Replicates | Calculation | Requirement | V1M6 1.7.3.3 b) ii. | The laboratory shall document the calculation for relative percent difference or other statistical treatments.
1713 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 b) Replicates | Evaluation | Requirement | V1M6 1.7.3.3 b) iii. | Results are compared to the acceptance criteria as published in the mandated method.
1714 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 b) Replicates | Evaluation | Requirement | V1M6 1.7.3.3 b) iii. | Where there are no established criteria, the laboratory shall determine internal criteria and document the method used to establish the limits.
1715 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 b) Replicates | Corrective Action | Requirement | V1M6 1.7.3.3 b) iii. | For replicate results outside established criteria, corrective action shall be documented or the data reported with appropriate data qualifying codes.
1716 | 1.7.3 Data Acceptance / Rejection Criteria | 1.7.3.3 b) Replicates | Reporting | Requirement | V1M6 1.7.3.3 b) iv. | The occurrence of a failed replicate and any associated actions shall be noted in the laboratory report to the client.
1717 | 1.7.4 Sample Handling | 1.7.4 Sample Handling |  |  | V1M6 | Sample Handling
1718 | 1.7.4 Sample Handling | 1.7.4 Sample Handling | Thermal Preservation | Requirement | V1M6 1.7.4 a) | All samples that require thermal preservation shall be considered acceptable if the arrival temperature of a representative sample container is either within 2°C of the required temperature or the method specified range.
1719 | 1.7.4 Sample Handling | 1.7.4 Sample Handling | Thermal Preservation | Requirement | V1M6 1.7.4 a) | For samples with a specified temperature of 4°C, samples with a temperature ranging from just above the freezing temperature of water to 6°C shall be acceptable.
1720 | 1.7.4 Sample Handling | 1.7.4 Sample Handling | Thermal Preservation | Exception | V1M6 1.7.4 a) i. | Samples that are delivered to the laboratory on the same day they are collected may not meet the requirements of Section 1.7.4.a. In these cases, the samples shall be considered acceptable if the samples were received on ice.
1721 | 1.7.4 Sample Handling | 1.7.4 Sample Handling | Thermal Preservation | Exception | V1M6 1.7.4 a) ii. | If sample analysis is begun within fifteen (15) minutes of collection, thermal preservation is not required. Thermal preservation is not required in the field if the laboratory receives and refrigerates the sample within fifteen (15) minutes of collection.
1722 | 1.7.4 Sample Handling | 1.7.4 Sample Handling | Chemical Preservation | Procedure | V1M6 1.7.4 b) | The laboratory shall implement procedures for checking chemical preservation using readily available techniques, such as pH or chlorine, prior to or during sample preparation or analysis.

## Module 7
ID | Major Topic | Topic | Subject | Category | Citation | Does the laboratory comply with this section? | Yes | No | N/A
1723 |  |  |  |  | V1M7 | TOXICITY TESTING
1724 | 1.1 Introduction |  |  |  | V1M7 1 | Introduction
1725 | 1.1 Introduction |  |  |  | V1M7 1.1 | This Standard applies to laboratories measuring the toxicity and/or bioaccumulation of contaminants in effluents (whole effluent toxicity or WET), receiving waters, sediments, elutriates, leachates and soils. In addition to the essential quality control standards described below, some methods may have additional or other requirements based on factors such as the type of organism evaluated and contain detailed quality control requirements for toxicity testing activities. The evaluation of laboratories for this discipline is in conjunction with a quality system as specified in the general requirements module. Adherence to quality systems requirements will ensure that all quality control procedures specified in this module are being followed.
1726 | 1.2 Scope |  |  |  | V1M7 1.2 | Scope
1727 | 1.2 Scope |  |  |  | V1M7 1.2 | The essential quality control procedures applicable to toxicity measurements are included in this Standard.
1728 | 1.2 Scope |  |  |  | V1M7 1.2 | Additional quality control requirements that are specified by method, regulation or project shall be met by laboratories.
1729 | 1.3 Terms and Definitions |  |  |  | V1M7 1.3 | Terms and Definitions
1733 | 1.3 Terms and Definitions |  |  |  | V1M7 1.3.2 | Exclusions and Exceptions
1731 | 1.3 Terms and Definitions |  |  |  | V1M7 1.3.1 | Additional Terms and Definitions
1732 | 1.3 Terms and Definitions |  |  |  | V1M7 1.3.1 | When referred to in this module, “sensitivity” relates to the meaning referenced in the accredited method.
1730 | 1.3 Terms and Definitions |  |  |  | V1M7 1.3 | The relevant definitions from TNI, Volume 1, Module 2, Section 3.0 apply. Definitions related to this document, which are used differently or do not exist in the above references are defined below.
1734 | 1.4 Method Selection |  |  |  | V1M7 1.3.2 | Reserved
1735 | 1.4 Method Selection |  |  |  | V1M7 1.4 | Method Selection
1736 | 1.4 Method Selection | Non Reference Method | Customer Communication | Requirement | V1M7 1.4 | When it is necessary to use testing methods not covered by an approved method, these shall be subject to agreement with the data user and shall include a clear specification of the data user’s requirements and the purpose of the environmental test.
1737 | 1.4 Method Selection | Non Reference Method | Validation | Requirement | V1M7 1.4 | The method developed shall have been validated appropriately before use.
1738 | 1.4 Method Selection | Non Reference Method | Assessment | Requirement | V1M7 1.4 | The characteristics of validated methods (e.g., the uncertainty of the results, limit of repeatability and/or reproducibility, robustness against external influences and/or cross-sensitivity against interference from the matrix of the sample/test object), as assessed for the intended use, shall be relevant to the users’ needs.
1739 | 1.5 Method Validation |  |  |  | V1M7 1.5 | Method Validation
1740 | 1.5 Method Validation |  |  |  | V1M7 1.5 | Validation is the confirmation by examination and the objective evidence that the particular requirements for a specific intended use are fulfilled.
1741 | 1.6 DOC |  |  |  | V1M7 1.6 | Demonstration of Capability (DOC)
1742 | 1.6 DOC | 1.6.1 General |  |  | V1M7 1.6.1 | General
1743 | 1.6 DOC | 1.6.1 General |  |  | V1M7 1.6.1 | Prior to acceptance and institution of any method for data reporting, satisfactory initial DOC is required (see Section 1.6.2).
1744 | 1.6 DOC | 1.6.1 General | Ongoing DOC | Requirement | V1M7 1.6.1 | Thereafter, ongoing DOC (Section 1.6.3), as per the quality control requirements in Section 1.7.1.2 is required.
1745 | 1.6 DOC | 1.6.1 General | DOC | Exception | V1M7 1.6.1 | In cases where a laboratory analyzes samples using a method that has been in use by the laboratory for at least one year prior to applying for accreditation, and there have been no significant changes in personnel or method, the ongoing DOC shall be acceptable as an initial DOC. The laboratory shall have records on file to demonstrate that an initial DOC is not required.
1746 | 1.6 DOC | 1.6.1 General |  |  | V1M7 1.6.1 | For the initial DOC, appropriate records as discussed in Section 1.6.2.1 shall be completed. (1.6.2.1)
1747 | 1.6 DOC | 1.6.1 General |  |  | V1M7 1.6.1 | An initial DOC shall be completed each time there is a change in personnel, or method.(1.6.2)
1748 | 1.6 DOC | 1.6.1 General | DOC Performance | Option | V1M7 1.6.1 | In general, this demonstration does not test the performance of the method in real world samples. However, before any results are reported, the initial DOC shall be performed. An initial DOC may be completed by a group of analysts and is for situations in which several individuals perform part of a set of activities that would produce a testing result.
1749 | 1.6 DOC | 1.6.1 General | Documentation | Requirement | V1M7 1.6.1 | All demonstrations shall be documented. All data applicable to the demonstration shall be retained and readily available at the laboratory.
1750 | 1.6 DOC | 1.6.2 Initial DOC |  |  | V1M7 1.6.2 | Initial DOC
1751 | 1.6 DOC | 1.6.2 Initial DOC | Frequency | Requirement | V1M7 1.6.2 | An initial DOC shall be made prior to using any method.
1752 | 1.6 DOC | 1.6.2 Initial DOC | Frequency | Requirement | V1M7 1.6.2 | An initial DOC shall be made at any time there is a significant change in personnel or method or any time that a method has not been performed by the laboratory or analyst in a twelve (12) month period.
1753 | 1.6 DOC | 1.6.2 Initial DOC | Documentation | Record | V1M7 1.6.2.1 | The laboratory shall document each initial DOC in a manner such that the following information is available for each affected employee:
1754 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M7 1.6.2.1 a) | analyst(s) involved in preparation and/or analysis;
1755 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M7 1.6.2.1 b) | matrix;
1756 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M7 1.6.2.1 c) | species and endpoint(s);
1757 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M7 1.6.2.1 d) | identification of method(s) performed;
1758 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M7 1.6.2.1 e) | identification of laboratory-specific SOP used for analysis, including revision number;
1759 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M7 1.6.2.1 f) | date(s) of analysis;
1760 | 1.6 DOC | 1.6.2 Initial DOC | DOC Record | Contents | V1M7 1.6.2.1 g) | summary of analyses, including information outlined in Section 1.6.2.2.
1761 | 1.6 DOC | 1.6.2 Initial DOC | Procedure | Requirement | V1M7 1.6.2.2 | If the method or regulation does not specify an initial DOC, the following procedure is acceptable. It is the responsibility of the laboratory to document that other approaches to initial DOC are adequate.
1762 | 1.6 DOC | 1.6.2 Initial DOC | DOC Procedure | Option | V1M7 1.6.2.2 | Each analyst shall meet the quality control requirements as specified in Section 1.7.1.2.
1763 | 1.6 DOC | 1.6.3 Ongoing DOC |  |  | V1M7 1.6.3 | Ongoing DOC
1764 | 1.6 DOC | 1.6.3 Ongoing DOC | Documentation | Procedure | V1M7 1.6.3 | The laboratory shall have a documented procedure describing ongoing DOC.
1765 | 1.6 DOC | 1.6.3 Ongoing DOC | Analyst | Requirement | V1M7 1.6.3 | The analyst(s) shall demonstrate on-going capability by meeting the quality control requirements of the method, laboratory SOP, client specifications, and/or this Standard.
1766 | 1.6 DOC | 1.6.3 Ongoing DOC | Alternate Procedures | Requirement | V1M7 1.6.3 | It is the responsibility of the laboratory to document that other approaches to on-going demonstration of capability are adequate.
1767 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M7 1.6.3 | This on-going demonstration may include performing another initial demonstration of capability as per 1.6.2.
1768 | 1.6 DOC | 1.6.3 Ongoing DOC | Procedure | Option | V1M7 1.6.3 | A documented process of analyst review using QC samples can serve as the annual on-going demonstration of capability. QC samples shall be reviewed to identify patterns for individuals or groups of analysts and determine if corrective action or retraining is necessary.
1769 | 1.7 Technical Requirements |  |  |  | V1M7 1.7 | Technical Requirements
1770 | 1.7 Technical Requirements | 1.7.1 Quality Control |  |  | V1M7 1.7.1 | Quality Control
1771 | 1.7 Technical Requirements | 1.7.1 Quality Control |  |  | V1M7 1.7.1 | The laboratory shall have quality control procedures for monitoring the validity of environmental tests undertaken. (see requirements listed in 1.7.1)
1772 | 1.7 Technical Requirements | 1.7.1 Quality Control | Result Review | Requirement | V1M7 1.7.1 | The resulting data shall be recorded in such a way that trends are detectable and, where practicable, statistical techniques shall be applied to the reviewing of the results.
1773 | 1.7 Technical Requirements | 1.7.1 Quality Control | Planned Review | Requirement | V1M7 1.7.1 | This monitoring shall be planned and reviewed and may include, but not be limited to, the following:
1774 | 1.7 Technical Requirements | 1.7.1 Quality Control | Planned Review | Option | V1M7 1.7.1 b) | participation in inter-laboratory comparison or proficiency-testing program;
1775 | 1.7 Technical Requirements | 1.7.1 Quality Control | Planned Review | Option | V1M7 1.7.1 c) | replicate tests using the same or different methods;
1776 | 1.7 Technical Requirements | 1.7.1 Quality Control | Planned Review | Option | V1M7 1.7.1 d) | retesting of retained samples; and
1777 | 1.7 Technical Requirements | 1.7.1 Quality Control | Planned Review | Option | V1M7 1.7.1 e) | correlation of results for different characteristics of a sample (for example, total phosphate should be greater than or equal to orthophosphate).
1778 | 1.7 Technical Requirements | 1.7.1.1 Essential QC |  |  | V1M7 1.7.1.1 | Essential Quality Control Procedures
1779 | 1.7 Technical Requirements | 1.7.1.1 Essential QC |  |  | V1M7 1.7.1.1 | These general quality control principles shall apply, where applicable, to all testing laboratories. The manner in which they are implemented is dependent on the types of tests performed by the laboratory and are further described in this module.
1780 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Methods | Requirement | V1M7 1.7.1.1 | The standards for any given test type shall assure that the applicable principles are addressed:
1781 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Quality Control | Procedure | V1M7 1.7.1.1 a) | All laboratories shall have detailed written protocols in place to monitor the following quality controls:
1782 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Positive and Negative Controls | Topic | V1M7 1.7.1.1 a) i) | positive and negative controls to monitor tests such as blanks, spikes, reference toxicants;
1783 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Variability | Topic | V1M7 1.7.1.1 a) ii) | tests to define the variability and/or repeatability of the laboratory results such as replicates;
1784 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Method Capability | Topic | V1M7 1.7.1.1 a) iii) | measures to evaluate method capability, such as percent minimum significant difference (PMSD);
1785 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Formulae Selection | Topic | V1M7 1.7.1.1 a) iv) | selection of appropriate formulae to reduce raw data to final results such as regression and statistical analyses;
1786 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Reagent and Standard Selection | Topic | V1M7 1.7.1.1 a) v) | selection and use of reagents and standards of appropriate quality;
1787 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Selectivity | Topic | V1M7 1.7.1.1 a) vi) | measures to assure the selectivity of the test for its intended purpose; and
1788 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Constant and Consistent Test Conditions | Topic | V1M7 1.7.1.1 a) vii) | measures to assure constant and consistent test conditions (both instrumental and environmental) where required by the method such as temperature, humidity, light or specific equipment conditions.
1789 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Evaluation | Requirement | V1M7 1.7.1.1 b) | All quality control measures shall be assessed and evaluated on an ongoing basis, and quality control acceptance criteria shall be used to determine the usability of the data.
1790 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Acceptance Criteria | Procedure | V1M7 1.7.1.1 c) | The laboratory shall have procedures for the development of acceptance/rejection criteria where no method or regulatory criteria exist.
1791 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | QC Procedures | Requirement | V1M7 1.7.1.1 d) | The quality control protocols specified by the laboratory’s method manual shall be followed.
1792 | 1.7 Technical Requirements | 1.7.1.1 Essential QC | Different QC Procedures | Requirement | V1M7 1.7.1.1 d) | The laboratory shall ensure that the essential standards outlined in this document or regulations (whichever are more stringent) are incorporated into their method manuals. When it is not apparent which is more stringent, the QC in the regulations is to be followed.
1793 | 1.7 Technical Requirements | 1.7.1.2 Positive and Negative Controls |  |  | 1.7.1.2 | Positive and Negative Controls
1794 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Reference Toxicants |  | V1M7 1.7.1.2 a) | Positive Control. Reference toxicant tests demonstrate a laboratory's ability to obtain consistent results with the method and evaluate the overall health and sensitivity of test organisms over time.
1795 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Reference Toxicants | Requirement | V1M7 1.7.1.2 a) i) | The laboratory shall demonstrate its ability to obtain consistent results with standard reference toxicants (SRT).
1796 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Frequency | Requirement | V1M7 1.7.1.2 a) ii) | Ongoing laboratory performance shall be demonstrated by performing routine SRT testing for each method, species and endpoint in accordance with the minimum frequency requirements specified in Section 1.7.1.2.a)iii).
1797 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Frequency | Requirement | V1M7 1.7.1.2 a) iii) | The frequency of ongoing laboratory reference toxicant testing shall be as follows unless the method specifically requires less frequent SRT tests (e.g., sediment tests).
1798 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Tests Performed Monthly | Requirement | V1M7 1.7.1.2 a) iii) | For methods conducted at a frequency of monthly or greater, SRT tests shall be conducted monthly.
1799 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Tests Performed > 1 Month | Requirement | V1M7 1.7.1.2 a) iii) | For methods and species commonly used in the laboratory, but which are tested at a frequency of less than monthly, SRT tests shall be conducted concurrently with the environmental test.
1800 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Outside Source Sensitivity | Requirement | V1M7 1.7.1.2 a) iii) | If the test organisms are obtained from an outside source, the sensitivity of each batch of organisms received from a supplier shall be determined via a concurrent SRT test
1801 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Outside Source Sensitivity | Exception | V1M7 1.7.1.2 a) iii) | If the test organisms are obtained from an outside source, the sensitivity of each batch of organisms received from a supplier shall be determined via a concurrent SRT test unless the supplier can provide control chart data for the last five SRT tests using the same SRT and test conditions. Supplied SRT data may not be older than six (6) months.
1802 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Specific SRT |  | V1M7 1.7.1.2 a) iv) | These standards do not currently specify a particular reference toxicant and dilution series.
1803 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Specific SRT | Requirement | V1M7 1.7.1.2 a) iv) | However, if the regulation identifies a reference toxicant or dilution series for a particular test, the laboratory shall follow the specified requirements.
1804 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | SRT in Test | Requirement | V1M7 1.7.1.2 a) iv) | All reference toxicant tests conducted for a given method and species shall use the same reference toxicant, test concentrations, dilution water and data analysis methods.
1805 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | SRT Dilution Factor | Requirement | V1M7 1.7.1.2 a) iv) | A dilution factor of 0.5x or greater shall be used for both acute and chronic tests.
1806 | 1.7 Technical Requirements | 1.7.1.2 a) Positive Control | Method Requirements | Requirement | V1M7 1.7.1.2 a) v) | The reference toxicant tests shall be conducted following the procedures required in the method.
1807 | 1.7 Technical Requirements | 1.7.1.2 b) Negative Controls |  |  | V1M7 1.7.1.2 b) | Negative Controls – Control, Brine Control, Control Sediment, Control Soil or Dilution Water
1808 | 1.7 Technical Requirements | 1.7.1.2 b) Negative Controls | Method Negative Controls | Requirement | V1M7 1.7.1.2 b) i) | The standards for the use, type and frequency of testing of negative controls are specified by the methods and by permit or regulation and shall be followed.
1809 | 1.7 Technical Requirements | 1.7.1.2 b) Negative Controls | Frequency | Requirement | V1M7 1.7.1.2 b) i) | A negative control is included with each test to evaluate test performance and the health and sensitivity of the specific batch of organisms.
1810 | 1.7 Technical Requirements | 1.7.1.2 b) Negative Controls | Additional Controls | Option | V1M7 1.7.1.2 b) ii) | Appropriate additional negative controls shall be included when sample adjustments (for example addition of thiosulfate for dechlorination) or solvent carriers are used in the test.
1811 | 1.7 Technical Requirements | 1.7.1.3 Variability |  |  | V1M7 1.7.1.3 | Variability and/or Reproducibility
1812 | 1.7 Technical Requirements | 1.7.1.3 Variability | Precision | Requirement | V1M7 1.7.1.3 | Intra-laboratory precision shall be determined on an ongoing basis through the use of further reference toxicant tests and related control charts as described above.
1813 | 1.7 Technical Requirements | 1.7.1.4 Test Sensitivity |  |  | V1M7 1.7.1.4 | Test Sensitivity
1814 | 1.7 Technical Requirements | 1.7.1.4 Test Sensitivity | PMSD | Requirement | V1M7 1.7.1.4 a) | The PMSD shall be calculated according to the formula specified by the method and reported with the test results.
1815 | 1.7 Technical Requirements | 1.7.1.4 Test Sensitivity | Point Estimates | Requirement | V1M7 1.7.1.4 b) | Point estimates: (LCp, ICp, or ECp) – Confidence intervals shall be reported as a measure of the precision around the point estimate value, when the calculation is possible.
1816 | 1.7 Technical Requirements | 1.7.1.5 Reagents and Standards |  |  | V1M7 1.7.1.5 | Selection and Use of Reagents and Standards
1817 | 1.7 Technical Requirements | 1.7.1.5 Reagents and Standards | Grade of Reagents | Requirement | V1M7 1.7.1.5 a) | The grade of all reagents used in toxicity tests is specified in the method except the reference standard.
1818 | 1.7 Technical Requirements | 1.7.1.5 Reagents and Standards | SRT Grade |  | V1M7 1.7.1.5 a) | All reference standards shall be prepared from chemicals that are analytical reagent grade or better.
1819 | 1.7 Technical Requirements | 1.7.1.5 Reagents and Standards | Preparation | Record | V1M7 1.7.1.5 a) | The preparation of all standards and reference toxicants shall be documented.
1820 | 1.7 Technical Requirements | 1.7.1.5 Reagents and Standards | Chemical Tests | Requirement | V1M7 1.7.1.5 b) | All standards and reagents associated with chemical measurements, such as dissolved oxygen, pH or specific conductance, shall comply with the Chemistry Module.
1821 | 1.7 Technical Requirements | 1.7.1.5 Reagents and Standards | Water | Requirement | V1M7 1.7.1.5 c) | Only reagent-grade water collected from distillation or de-ionization units is used to prepare reagents.
1822 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions |  |  | V1M7 1.7.1.6 | Constant and Consistent Test Conditions
1823 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Separation of Tasks | Requirement | V1M7 1.7.1.6 a) | If closed refrigerator-sized incubators are used, culturing and testing of organisms shall be separated to avoid cross-contamination.
1824 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Laboratory Facilities | Requirement | V1M7 1.7.1.6 b) | Laboratory space shall be adequate for the types and numbers of tests performed. The building shall provide adequate cooling, heating and illumination for conducting testing and culturing; hot and cold running water shall be available for cleaning equipment.
1825 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Air Quality | Requirement | V1M7 1.7.1.6 c) | Air used for aeration of test solutions, dilution waters and cultures shall be free of oil and fumes.
1826 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Taxonomic ID | Requirement | V1M7 1.7.1.6 d) | The laboratory or a contracted outside expert shall positively identify test organisms to species on an annual basis.
1827 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Taxonomic ID | Record | V1M7 1.7.1.6 d) | The taxonomic reference (citation and page(s)) and the names(s) of the taxonomic expert(s) shall be kept on file at the laboratory.
1828 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Outside Source ID | Requirement | V1M7 1.7.1.6 d) | When organisms are obtained from an outside source the supplier shall provide this same information.
1829 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Chemical and Physical Tests | Requirement | V1M7 1.7.1.6 e) | Equipment used for routine support measurements of chemical and physical parameters such as pH, DO, conductivity, salinity, alkalinity, hardness, chlorine, ammonia and weight shall be calibrated, and/or standardized per manufacturer’s instructions.
1830 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Results and Calibration | Record | V1M7 1.7.1.6 e) | All measurements and calibrations shall be documented.
1831 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Temperature | Requirement | V1M7 1.7.1.6 f) | Test temperature shall be maintained as specified for the method.
1832 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Temperature control equipment | Requirement | V1M7 1.7.1.6 f) | Temperature control equipment shall be adequate to maintain the required test temperature(s).
1833 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Average Daily Temperature | Requirement | V1M7 1.7.1.6 f) | The average daily temperature of the test solutions shall be maintained within method specified range.
1834 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Frequency | Requirement | V1M7 1.7.1.6 f) | The minimum frequency of measurement shall be once per twenty-four (24) hour period. The test temperature for continuous-flow toxicity tests shall be recorded and monitored continuously.
1835 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Continuous Loggers | Requirement | V1M7 1.7.1.6 f) | Where electronic data loggers are used, temperature shall be monitored at a frequency sufficient to capture temporal variations of the environmental control system.
1836 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Water | Requirement | V1M7 1.7.1.6 g) | Reagent grade water, prepared by any combination of distillation, reverse osmosis, ion exchange, activated carbon and particle filtration, shall meet the method specified requirements.
1837 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Dilution Water | Requirement | V1M7 1.7.1.6 h) | The quality of the standard dilution water used for testing or culturing shall be sufficient to allow satisfactory survival, growth and reproduction of the test species as demonstrated by routine reference toxicant tests and negative control performance.
1838 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Water Testing | Requirement | V1M7 1.7.1.6 h) | Water used for culturing and testing shall be analyzed for toxic metals and organics whenever the minimum acceptability criteria for control survival, growth or reproduction are not met and no other cause, such as contaminated glassware or poor stock, can be identified.
1839 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Food | Requirement | V1M7 1.7.1.6 i) | The quality of the food used for testing or culturing shall be sufficient to allow satisfactory survival, growth and reproduction of the test species as demonstrated by routine reference toxicant tests and negative control performance.
1840 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Food Evaluation | Procedure | V1M7 1.7.1.6 i) | The quality of the food used for testing or culturing shall be sufficient to allow satisfactory survival, growth and reproduction of the test species as demonstrated by routine reference toxicant tests and negative control performance. The laboratory shall have written procedures for the evaluation of food acceptance.
1841 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Bioaccumulation | Requirement | V1M7 1.7.1.6 j) | A subset of organisms used in bioaccumulation tests shall be analyzed at the start of the test (baseline) for the target compounds to be measured in the bioaccumulation tests.
1842 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Chamber | Requirement | V1M7 1.7.1.6 k) | Test chamber size and test solution volume shall be as specified in the method.
1843 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Chamber | Requirement | V1M7 1.7.1.6 k) | Test chamber size and test solution volume shall be as specified in the method. All test chambers used in a test shall be identical.
1844 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Feeding Regimen | Requirement | V1M7 1.7.1.6 l) | Test organisms shall be fed the quantity and type food or nutrients specified in the method. T
1845 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Feeding Frequency | Requirement | V1M7 1.7.1.6 m) | All organisms in a test shall be from the same source and lot. Where available, certified seeds are used for soil tests.
1846 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Organism Health | Requirement | V1M7 1.7.1.6 n) | All organisms used in tests, or used as broodstock to produce neonate test organisms (for example cladocerans and larval fish), shall appear healthy, show no signs of stress or disease and exhibit acceptable survival (90% or greater) during the twenty-four (24) hour period immediately preceding use in tests.
1847 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Equipment Construction | Requirement | V1M7 1.7.1.6 o) | All materials used for test chambers, culture tanks, tubing, etc. and coming in contact with test samples, solutions, control water, sediment or soil or food shall be non-toxic and cleaned as described in the methods. Materials shall not reduce or add to sample toxicity. Appropriate materials for use in toxicity testing and culturing are described in the methods.
1848 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Light Intensity | Requirement | V1M7 1.7.1.6 p) | Light intensity shall be maintained as specified in the methods.
1849 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Light Intensity Measurements | Requirement | V1M7 1.7.1.6 p) | Measurements shall be made and recorded on a yearly basis.
1850 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Photoperiod | Requirement | V1M7 1.7.1.6 p) | Photoperiod shall be maintained as specified in the methods
1851 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Photoperiod Frequency | Requirement | V1M7 1.7.1.6 p) | Photoperiod shall be shall be documented at least quarterly.
1852 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Light Intensity Measurements | Requirement | V1M7 1.7.1.6 p) | For algal and plant tests, the light intensity shall be measured and recorded at the start of each test.
1853 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Organism Health and Culturing | Record | V1M7 1.7.1.6 q) | The health and culturing conditions of all organisms used for testing shall be documented by the testing laboratory. Such documentation shall include culture conditions (e.g. salinity, hardness, temperature, pH) and observations of any stress, disease or mortality.
1854 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Organism Health and Culturing Records | Contents | V1M7 1.7.1.6 q) | The health and culturing conditions of all organisms used for testing shall be documented by the testing laboratory. Such documentation shall include culture conditions (e.g. salinity, hardness, temperature, pH) and observations of any stress, disease or mortality.
1855 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Specification for outside sources | Requirement | V1M7 1.7.1.6 q) | When organisms are obtained from an outside source, the laboratory shall obtain written documentation of these water quality parameters and biological observations for each lot of organism received. T
1856 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Specification for outside sources | Requirement | V1M7 1.7.1.6 q) | These observations shall adequately address the twenty-four (24) hour time period referenced in item 1.7.1.6 n) above.
1857 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Specification for outside sources | Record | V1M7 1.7.1.6 q) | The laboratory shall also record each of these observations and water quality parameters upon the arrival of the organisms at the testing laboratory.
1858 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Age/age range | Requirement | V1M7 1.7.1.6 r) | Age and the age range of the test organisms shall be as specified in the method.
1859 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Supporting Information | Record | V1M7 1.7.1.6 r) | Supporting information, such as hatch dates and times, times of brood releases and metrics (for example, chironomid head capsule width) shall be documented.
1860 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Holding time | Requirement | V1M7 1.7.1.6 s) | The maximum holding time of effluents (elapsed time from sample collection to first use in a test) shall not exceed thirty-six (36) hours; samples may be used for renewal up to seventy-two (72) hours after first use except as prescribed by the method and approved by the regulatory agency having authority for program oversight.
1861 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Replicates | Requirement | V1M7 1.7.1.6 t) | All tests shall have at least the minimum number of replicates per treatment as prescribed by the method.
1862 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Control Population Distribution | Requirement | V1M7 1.7.1.6 u) | The control population of Ceriodaphnia in chronic effluent or receiving water tests shall contain no more than 20% males.
1863 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | C dubia culturing | Requirement | V1M7 1.7.1.6 v) | The culturing of C. dubia shall be adequate such that blocking by parentage can be established.
1864 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | DO and pH | Requirement | V1M7 1.7.1.6 w) | Dissolved oxygen and pH in aquatic tests shall be within acceptable range at test initiation. Minimal aeration is provided to tests if acceptable dissolved oxygen concentrations cannot be otherwise maintained.
1865 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Tolerance range for solids | Requirement | V1M7 1.7.1.6 x) | Test soils or sediments shall be within the geochemical tolerance range of the test organism.
1866 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Unacceptable Test Conditions | Requirement | V1M7 1.7.1.6 y) | An individual test may be conditionally acceptable if temperature, dissolved oxygen, pH and other specified conditions fall outside specifications, depending on the degree of the departure and the objectives of the tests (see test conditions and test acceptability criteria specified for each method).
1867 | 1.7 Technical Requirements | 1.7.1.6 Constant & Consistent Test Conditions | Acceptability | Requirement | V1M7 1.7.1.6 y) | The acceptability of the test shall depend on the experience and professional judgment of the technical director and the permitting authority.
1868 | 1.7 Technical Requirements | 1.7.2 Data Acceptance / Rejection |  |  | V1M7 1.7.2 | Data Acceptance/Rejection Criteria
1869 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls |  |  | V1M7 1.7.2.1 | Positive Controls
1870 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | Control Performance and Endpoints | Record | V1M7 1.7.2.1 | A laboratory shall record the control performance and statistical endpoints (such as NOEC or ECp) for each method and species on control charts.
1871 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | Precisions Evaluation | Requirement | V1M7 1.7.2.1 | The laboratory shall also evaluate precision (i.e. coefficient of variation, CV) for these tests against method specific or laboratory-derived criteria to determine validity of the testing result.
1872 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | ICp, Ecp Criteria | Requirement | V1M7 1.7.2.1 | For endpoints that are point estimates (ICp, ECp), control charts are constructed by plotting the cumulative mean and the control limits, which consist of the upper and lower 95% confidence limits (+/- 2 standard deviations).
1873 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | NOEC, NOAEC Criteria | Requirement | V1M7 1.7.2.1 | For endpoints from hypothesis tests (NOEC, NOAEC) the values are plotted directly, and the control limits consist of one concentration interval above and below the concentration representing the central tendency (i.e. the mode).
1874 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | Point Estimates | Requirement | V1M7 1.7.2.1 | For endpoints that are point estimates the cumulative mean CV is calculated. These values are maintained on control charts.
1875 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | Hypothesis Tests | Requirement | V1M7 1.7.2.1 | For endpoints from hypothesis tests, the PMSD is calculated. These values are maintained on control charts.
1876 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | Control Chart Exceedances | Requirement | V1M7 1.7.2.1 | Control chart limits are expected to be exceeded occasionally regardless of how well a laboratory performs. Acceptance limits for point estimates (ICp, ECp) that are based on 95% confidence limits should theoretically be exceeded for one in twenty tests. Depending on the dilution factor and test sensitivity, control charts based on hypothesis test values (NOEC, NOAEC) may be expected to be exceeded on a similar frequency.
1877 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | Evaluation of Exceedances | Requirement | V1M7 1.7.2.1 | Test results that fall outside of control chart limits at a frequency of 5% or less, or which fall just outside control chart limits (especially in the case of highly proficient laboratories which may develop relatively narrow acceptance limits over time), are not rejected de facto. Such data are evaluated in comparison with control chart characteristics including the width of the acceptance limits and the degree of departure of the value from acceptance limits.
1878 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | Acceptance/ Rejection Criteria | Procedure | V1M7 1.7.2.1 | Laboratories shall develop acceptance/rejection policies, consistent with the methods, for SRT data which considers source of test organisms, the direction of the deviation, test dilution factor, test sensitivity (for hypothesis test values), testing frequency, out-of-control test frequency, relative width of acceptance limits, inter-test CV, and degree of difference between test results and acceptance limits.
1879 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | Corrective Action | Requirement | V1M7 1.7.2.1 | In the case of reference toxicant data which fails to meet control chart acceptance criteria, the test data are examined for defects, corrective action taken and the test repeated if necessary, using a different batch of organisms or the data is qualified.
1880 | 1.7 Technical Requirements | 1.7.2.1 Positive Controls | Intra-laboratory Precision | Requirement | V1M7 1.7.2.1 | Intra-laboratory precision is determined on an ongoing basis through the use of control charts. The control charts shall be plotted as point estimate values, such as EC25 for chronic tests and LC50 for acute tests, or as appropriate hypothesis test values, such as the NOEC or NOAEC, over time within a laboratory.
1881 | 1.7 Technical Requirements | 1.7.2.2 Negative Controls |  |  | V1M7 1.7.2.2 | Negative Controls
1882 | 1.7 Technical Requirements | 1.7.2.2 Negative Controls | Criteria | Requirement | V1M7 1.7.2.2 | The test acceptability criteria specified in the method shall be achieved for both the reference toxicant and the effluent or environmental sample toxicity test. The criteria shall be calculated and shall meet the method specified requirements for performing toxicity tests.
1883 | 1.7 Technical Requirements | 1.7.2.3 Statistical Method Selection |  |  | V1M7 1.7.2.3 | Selection of Appropriate Statistical Analysis Methods
1884 | 1.7 Technical Requirements | 1.7.2.3 Statistical Method Selection | Analysis Method | Requirement | V1M7 1.7.2.3 a) | Methods of data analysis and reporting as specified by language in the regulation, permit, or the method shall be followed as required.
1885 | 1.7 Technical Requirements | 1.7.2.3 Statistical Method Selection | Plotting Data | Requirement | V1M7 1.7.2.3 b) | Toxicity data shall be plotted on semi-logarithmic graph paper, relating time, mortality, and effluent concentration to verify computational results.
1886 | 1.7 Technical Requirements | 1.7.3 Sample Handling |  |  | V1M7 1.7.3 | Sample Handling
1887 | 1.7 Technical Requirements | 1.7.3 Sample Handling | Preservation | Requirement | V1M7 1.7.3 | All samples shall be chilled to 0-6°C during or immediately after collection except as prescribed by the method and approved by the regulatory agency having authority for program oversight.